In Compliance with the requirements of:

NEMA MEDPACS RFP -  to Develop Guidelines for Cross-Vendor DICOM Testing

PRESENTED by

Information and Control Technology, Inc.

and

Penn State University College of Medicine

Guidelines for Cross-Vendor DICOM Testing

Title:			Level 1 - Introduction and Overview

Document ID:		1.2.840.113702.1.1.1.4

Project Leader:		Franklyn T. Bradshaw, Ph.D.

Position Title:		President, Information and Control Technology

Mailing Address:	Information and Control Technology, Inc.

23600 Mercantile Rd., Suite 120

Beachwood, Ohio 44122

Telephone Number: 	(216) 831-9777

Project Period:  	7/24/95 - 7/1/96

Co-PI:			Fred W. Prior, Ph.D.

Chief, Section of Radiologic Computing and Imaging Science

Department of Radiology

Penn State College of Medicine,  Hershey, PA 17033

Document Version:  	6/4/96

INTRODUCTION

Interoperability is the fundamental reason for the existence of the DICOM standard.  For almost two decades the medical imaging community has searched for the Rosetta stone, the key facilitating technology, that will make painless multi-vendor connectivity a reality. The DICOM standard is intended to be that key.  Throughout three large-scale public demonstrations and numerous private efforts, it has been repeatedly shown that DICOM works.  Today, there are numerous imaging devices and systems on the market and installed in medical facilities that are capable of supporting DICOM. Currently there is no consistent  methodology for determining whether and to what extent two implementations actually interoperate. Perhaps more importantly, at least to the user community,  there is no means to determine if a given equipment combination has ever been tested. The initiative by the MEDPACS Section of NEMA to develop Guidelines (the Guidelines) for Cross-Vendor DICOM Testing, is an important first step toward a remedy for the DICOM interoperability testing dilemma. Progress in this direction requires a more complete refinement of concepts and definitions of terms. To be precise these Guidelines support testing of DICOM interconnectivity.

This document is the introductory document of a set of documents that present the Guidelines for Cross-Vendor DICOM Interconnectivity Testing initiated by the Medical Picture Archiving and Communications Section of the National Electrical Manufacturers Association (NEMA). The objective is to provide a set of Guidelines that vendors and other DICOM implementors can use in a consistent fashion to assure that their DICOM applications will interconnect and to provide users with a consistent view of DICOM product interconnectivity coverage.

DEFINITIONS AND SCOPE

Basic Network Interconnectivity

In the DICOM context, the basic network environment may be TCP/IP or ISO. Basic network connectivity in either environment is the ability to establish a connection and exchange messages at the basic network level.

DICOM Interconnectivity

DICOM interconnectivity addresses the ability of two Application Entities to establish a connection and to exchange DICOM messages in a manner consistent with the DICOM protocol.

DICOM Interconnectivity Testing

A formal process to validate that two implementations satisfy the requirements of DICOM Interconnectivity as defined herein.

DICOM Interoperability

Interoperability is the ability for two Application Entities to exchange Composite Information Objects or to exchange information about Normalized Information Objects. For Composite Information Objects, interoperability requires that an Information Object sent from one Application Entity will be completely received by the receiving Application Entity. For Normalized Information Objects, interoperability requires that the information sent from one Application Entity will have the intended effect on the Information Object maintained by the receiving Application Entity.

DICOM Interoperability Testing

A formal process to validate that two Application Entities satisfy the requirements of DICOM  Interoperability as defined herein.

The distinction between interconnectivity and interoperability can be reflected in a basic model of communication. (Figure 1.)

The basic concept of the Model of Communication is that there are abstract objects or pieces of information that must be communicated between the two communicating entities. The transfer mechanism of communication is not direct at the abstract level. In order for information to be communicated from one entity to another, the  sending entity must represent the information as symbols and transmit the symbols to the receiving entity. The receiving entity must then interpret the received symbols back into information objects and we can say that there is communication when the two communicating entities view and utilize the information objects in the same ways.

Communication in DICOM precisely matches the general model of communication presented above. In DICOM, information is contained in information objects, either as normalized information objects or as composite information objects. Information is moved from one application entity to another by one of two means. In the case of composite information objects, copies of the information object are represented in DICOM messages and  sent to the communicating  application entity. In the case of normalized information objects, information about the normalized information is represented in DICOM messages and sent to the communicating application entity. For both the sending and the receiving entity, the intent is that all of the information sent by the sending entity shall be received by the receiving entity and both entities should interpret the information in the same ways.

The goal of interconnectivity testing is to evaluate the extent to which two DICOM conformant devices can communicate.  Each device includes one or more DICOM application entities which implement specified (in the vendorÕs conformance statement) subsets of the DICOM standard.  At the most basic level application entities communicate with peer application entities. It is therefore appropriate that this and subsequent Guidelines documents focus at the level of the application entity rather than at the device level.  Any pair of devices under test may each support several application entities not all of which are complimentary (i.e. SCU-SCP pair for the same SOP class). Only complimentary application entity pairs can be considered for interconnectivity or DICOM interoperability testing.

Conformance Testing

A formal process to validate that an implementation behaves in accordance with its associated conformance statement with respect to the DICOM Standard.

Functional Interoperability

The ability to exchange medical information which has clinical utility on both sides of the connection.

In this document and related documents, we have substituted the term ‘Functional Interoperability’ for the term ‘interoperability’ as it is normally used. Normally, the intended notion of interoperability is not only to exchange information but to manipulate that information in both Application Entities with functions ranging from basic windowing and leveling to three-dimensional rendering. There are two main difficulties with attempting to establish a general set of guidelines for testing Functional Interoperability. The first difficulty is that the functions supported by each piece of equipment differ. The second difficulty is that all use of the information objects in the DICOM Information Model is outside of the DICOM Standard.

Scope of The Current Guidelines

The testing guidelines that have been prepared for cross-vendor testing address testing at the level of DICOM Interconnectivity.

Information

Information

Communication

Interpretation

Interpretation

Representation

Representation

Data

Data

Transmission

Figure 1.

Model of communication

GUIDELINES COMPONENTS

The IEEE has published a Testing Standard (IEEE Std. 829-1983) which lays the foundation for complete documentation of Test Plans and Procedures. The standard establishes document structure and testing procedure guidelines in the following areas:

- Test Plan

- Test Design Specification

- Test Case Specification

- Test Logging

- Test Procedures

- Test Summary Reporting

A more complete summary of  IEEE Std. 829-1983 is presented in Appendix A of this document. The IEEE standard should serve as the foundation for the components of the Guidelines for Cross-Vendor DICOM Testing.

GUIDELINES - DEVELOPMENT OF REQUIREMENTS

There are two major areas of the DICOM Standard which will act as the focus of Cross-Vendor Interconnectivity Testing:

1)	Associations (including Negotiables)

2)	Service Classes (Storage, Query/Retrieve, Print, Patient/Study/Results Management)

The Guidelines are intended to permit testing of DICOM implementations as specified in the following NEMA  Publications:

PS3.2-1993	Conformance

PS3.3-1993	Information Object Definitions

PS3.4-1993	Service Class Specifications

PS3.5-1993	Data Structures and Encoding

PS3.6-1993	Data Dictionary

PS3.7-1993	Message Exchange

PS3.8-1992	Network Communication Support for Message Exchange

All modifications to these base documents approved as of September 1, 1995 will also be considered. The Guidelines do not consider Supplements to the DICOM Standard.

VENDOR TESTERS - DEVELOPMENT STRATEGY

The MEDPACS section of NEMA plans to develop a reference list of vendors who have implemented the DICOM Standard to some level of conformance in equipment and who are willing to pair with other vendors to conduct interconnectivity testing. The point of contact within each registered organization will be published. The goal is to facilitate the process by which any vendor organization will identify another organization with which to test.

Joint development of test plans, test designs, and testing procedures creates a dilemma for vendors. The IEEE Standard recommends and design of the testing process requires access to internal design and implementation information about each system to be tested. The detailed internal information required to develop the testing process for a system is precisely the type of information that any vendor is reluctant to share with another organization that is ultimately a competitor.

The dilemma can be addressed by invoking some formal concepts of systems.

Open System

An open system is one that interacts with its external environment.

System - External View

The external view of a system addresses the set of interactions that the system is capable of having with its external environment.

System - Internal View

The internal view of a system addresses both the set of interactions that the system is capable of having with its external environment as well as the interactions among the component subsystems that make up the system.

These fundamental definitions provide sufficient basis to address the dilemma faced by organizations that wish to conduct connectivity testing with other cooperating organizations.

Each Application Entity that interacts with another Application Entity in conformance with the DICOM Standard is an open system. Each Application Entity presents an external view to another Application Entity and therefore to the testing team of another organization. Test plans and test designs for interconnectivity testing between two Application Entities will be based on the external views presented by both Application Entities.

The testing team of an organization may require design and implementation documents to create the internal view of their own Application Entity in order to fully define the external view and the testing process for that Application Entity. This process of using the detailed internal view of a system to define the external view is a step that will be performed only one time for each Application Entity. The detailed internal information of that Application Entity need not be shared with any external organization. Only the external view of the Application Entity needs to be available when defining the interconnectivity testing process with another Application Entity.

An open systems approach to joint interconnectivity testing of Application Entities suggests both a series of steps the vendors should take, individually and jointly, and also sets of documents to be prepared, both individually and jointly. The following material suggests both the documents to be prepared and the steps to be followed.

GUIDELINES - DEVELOPMENT STRATEGY

The strategy used to develop the Guidelines for Cross-Vendor DICOM Testing must accommodate the fact that there are a number of degrees of freedom within the statements of the DICOM standard and the fact that vendors may pair up in arbitrary ways to test. Some examples of the variations that are inherent in DICOM include:

1.	Service classes do not adhere to the same model .

2.	Service classes differ in the number of conformance levels that are acceptable within the standard.

The result of all of the variability is that it is not practical to develop a single document that will serve as a guideline for all of the testing situations that will arise in DICOM interconnectivity. The approach proposed is to develop a hierarchy of document structures/templates that can be used to accommodate the wide range of testing situations that will be encountered in DICOM interconnectivity testing. The documents are organized in a four-level hierarchy as illustrated in Figure 2 and Figure 3.

Level 1 and Level 2 contain template documents. These template documents follow the outline defined by the IEEE standard. Level 1 documents include information which is common to all aspects of DICOM interconnectivity testing.

The second level template document sets are partitioned into categories of Association Services and Service Classes (Storage, Query/Retrieve, etc.). The goal at this level is to develop a set of template documents for each category that cover all aspects of testing within that category. Template  Test Plan documents present the overview of testing the categories of Association Services and Service Classes.

The third level of document sets will be produced by the test implementors as they adapt the Level 2 template documents to specific testing situations. Level 3 in the hierarchy represents the set of documents that a single vendor would prepare to test the Association Services and any of the Service Classes supported by an Application Entity. Level 3 in Figure 2 illustrates the document set produced by a vendor in the process of preparing, executing, and reporting the tests for the Storage Service Class and Query/Retrieve Service Class for one Application Entity.

Level 4, shown in Figure 3, represents the relationship and the essential information sharing between two vendors as they jointly design and execute the cross-vendor tests of their Application Entities.

Before the development of the guidelines and the template documents is complete, it will be necessary to develop a testing example that goes from the second level template document set to completed tests with completed test documentation. The conformance statements are essential components of the documentation that must be used for specific pairs of vendor’s equipment.

Table 1 summarizes the document set that is defined in IEEE Std. 829-1983 as well as the additional documents that are required for the overall structure defined for the DICOM interconnectivity testing. As indicated, documents that occur at more than one level become increasingly more detailed as the level increases.

Figure 3

Document Hierarchy

Level 3, Level 4

Joint Vendor Document Generation

Level 4

Vendor 1, Vendor 2

Conformance Statements

Cross Vendor Joint Document Set (Storage, etc.)

Level 3

Vendor 2

Conformance Statement

Vendor Specific Document Set (Storage, etc.)

Level 3

Vendor 1

Conformance Statement

Vendor Specific Document Set (Storage, etc.)

| Document Usage Table |  |  |  |  |
| --- | --- | --- | --- | --- |
| Document Name | Document Level |  |  |  |
|  | 1 | 2 | 3 | 4 |
| Test Plan | x | x | x | x |
| Test Design Specification | x | x | x | x |
| Test Case Specification | x | x | x | x |
| Test Log | x | x | x | x |
| Test Incident Report | x | x | x | x |
| Test Summary Report | x | x | x | x |
| Vendor’s Conformance Statement |  |  | x | x |
| Table 1. |  |  |  |  |

VENDOR TESTERS - DEVELOPMENT STEPS

1.	Test Preparation Steps

1.1	Individual Vendor Steps

(These steps will be taken one time for any Application Entity that a vendor expects to test for interconnectivity with Application Entities of other vendors.)

1.1.1	Review the template documents provided in the NEMA  document tree, the template documents at Level 2 and the testing example documents at Level 3.

1.1.2	For the Application Entity of interest, procure and review the necessary internal information required to develop the testing process documents pertinent to the Application Entity.

1.1.3	Prepare the testing process documents that will fully test the Application Entity for interconnectivity in conformance with the DICOM Standard: Test Plan, Test Design, Test Case documents.

1.2	Joint Vendors Steps

(These steps will be taken one time for each pair of Application Entities that vendors jointly agree to test.)

1.2.1	Jointly review the conformance statements and the testing process documents of the two Application Entities to be tested for interconnectivity. Determine the DICOM functions that the two Application Entities jointly support and the set of tests to be executed for interconnectivity testing of the two Entities.

1.2.2	Jointly prepare the testing process documents required to test interconnectivity between the two Application Entities.

The set of DICOM functions that are jointly supported by the two Application Entities is the intersection of the set supported by each Application Entity.

2.	Test Execution and Reporting Steps

2.1	Individual Vendor Steps

2.1.1	Log each test and test result for the vendor Application Entity, recording incident (unexpected result) and whether the incident was resolved during the testing.

2.1.2	Prepare the set of test result documents - Test Log, Test Incident Report, Test Summary Report.

(These documents pertain only to this vendor’s Application Entity.)

2.2	Joint Vendor Steps

(These steps require mutual consent between the two vendor teams.)

2.2.1	Prepare a joint Test Summary Report, summarizing the joint test results of interconnectivity testing between the two Application Entities.

2.2.2	Publish the joint Test Plan and the joint Test Summary Report through NEMA in the tree of documents defined in this report.

GUIDELINES DOCUMENTS  -  UNIQUE IDENTIFICATION

All template and example documents at all levels of the hierarchy shall be uniquely identified by a Unique Identifier (UID) formatted according to NEMA Publication PS3.5-1993. The UID prefix for any document shall be defined by the parent organization. Currently, the Organization Root for the Guidelines documents is Penn State University which has defined a UID prefix: 1.2.840.113702.1. for all Guidelines documents. Ultimately, NEMA will become the Organization Root for the Guidelines documents.

The UID suffix for Guidelines documents and for documents submitted by vendors shall be created as follows:

.<hierarchy level>.<document number>.<version number>

The UID of this document (version 1) is : 1.2.840.113702.1.1..1.

Test Implementor generated documents based on the templates shall be identified in an analogous manner. One member of the testing pair shall be designated by mutual agreement to provide the root (organization ID) for each testing instance.

DOCUMENTS MANAGEMENT AND PUBLICATION

The World-Wide Web (WWW or, the Web) is the most rapidly growing segment of the Internet. At Penn State University, a Web server has been in operation since early 1994 (URL: http://www.xray.hmc.psu.edu). The Penn State anonymous ftp and Web servers have been extended to support electronic publication of the Guidelines (URL: http://www.xray.hmc.psu.edu/dicom/guidelines/Guidelines.html). As the Guidelines documents stabilize, the home page for the documents will be moved to NEMA for on-going maintenance.

The Guidelines authors have prepared a complete set of sample documents including, the Level 1 base documents, Level 2 template documents and an example document set at Level 3, Level 4 from a specific pair of systems. Vendors performing interconnectivity tests using these Guidelines are requested to submit the joint Test Plan and joint Test Summary Report for each pair of systems tested. All testing documents will be written in html. Vendors may submit their completed test documents via anonymous ftp to the Penn State server for publication (ftp.xray.hmc.psu.edu). An e-mail distribution list will be established so that all test contacts may communicate with each other by sending a single e-mail.  The vendor contact list will be maintained on the ftp server and at NEMA’s discretion published on the Web. Test results will be compiled and published on the Web. Storage will also be provided on an  ftp server for vendor conformance statements and these will be published in html versions on the Web at the vendor’s discretion. As indicated, the ftp and Web server resources reside at Penn State but will be moved to NEMA for on-going maintenance as the documents stabilize.

INTERCONNECTIVITY TESTING CONSTRAINTS

Several factors impact test design for interconnectivity testing of Application Entities. The first restriction is that the testing guidelines should not require invasive change in the Application Entities to be tested. On the other hand, the testability of any system is determined by the test points that have been made accessible by the system designers and implementors, both for injecting and extracting data. In these guidelines the premise is adopted that system designers and implementors have made or will make available an adequate amount of information for testing purposes. The actual level of information accessible for testing any pair of systems will depend on the decisions made at design and implementation time. The test plan that is generated and the tests that are designed for a given pair of Application Entities must be consistent with and adhere to the operational environment provided by the designers and implementors of the Application Entities.

The overall test plan and the design of tests will depend on the amount of information and the control that the designers and implementors have made available to the system operators. In testing for interconnectivity of systems that conform to the DICOM Standard, there are several levels of information access that may be available to the user:

Event Access:	Access to the information that a significant event has occurred.

Command Data Access:	Access to the relevant portions of the command set information.

Data Set Access:	Access to the information contained in the data set of the messages.

Within the DICOM operational context, a user is not normally aware of the sequences of events associated with interactions between systems. At the user level, an action may be requested which implies the initiation of a DICOM message transfer. Normally, the only indication that the user has that the action has taken place, either correctly or incorrectly, is through the log files that are provided in the implementation.

User Access Points:

Input:

User interface

Output:

User log file

Log files for:

ACSE Association Service Provider

DIMSE Service Provider

Application Entity Log file of ACSE and DIMSE services called.

The significant stages that take place in interconnection are:

Establishing an Association

Invoking DIMSE Services over the Association

Releasing the Association

Aborting an Association (abnormally)

Two types of information can be moved between Application Entities - image data and text data. In more conventional testing situations, testing for the integrity of data transfer or some transformation on the data is more straight-forward. The data at the beginning and end of the operational process is available and it can be examined by alternative means to validate that the proper transfer or transformation has taken place. In testing for DICOM interconnectivity, these conventional circumstances for testing do not hold. The data at each end may be available, in its own environment, in an indirect way at the user interface. In the normal DICOM interconnection context, there is no good way to ‘close the loop’ on data transfers and transformations.

The difficulty of trying to apply conventional test strategies to the DICOM interconnection situation is that the data may change format several times in going from one inter-operating system to another. The difficulty with the format change in the DICOM environment is that neither Application Entity  knows the internal data format of the inter-operating Application  Entity.

In conventional testing, the testing group has both the ability to tell if and when an event has occurred and has access to the data at critical points in the data transfer or transformation process to determine that the process is proceeding in the expected way. In the testing for DICOM interconnectivity, the normal ability to test at critical points along the process path may not be available.

A test for interconnectivity between two Application Entities has the objective to show that the Application Entities are able to exchange messages in accordance with the DICOM Standard. Successful exchange of messages does not guarantee that clinically useful transfer of information will result and the real objective of message exchange is to achieve clinically useful information transfer.

On the other hand, a test for interconnectivity between two Application Entities is not intended to be a detailed test process for debugging either or both of the Application Entities involved. Testing should be kept at a level that will demonstrate successful or unsuccessful interconnection without going down to a level to ‘debug’ why the two Application Entities fail to interconnect successfully. The scope of the test plan will include functionality testing of the Application Entities, the operator interfaces for invoking message transfer, and indications of successful and failed operations.

APPENDIX A:  IEEE Std 829-1983 Summary

IEEE Standard for Software Test Documentation

1.	Scope of the Standard

The standard defines and describes a set of basic test documents which are associated with the dynamic aspects of software testing, i.e. the execution of procedures and code. The standard defines the purpose, outline, and content of each basic document. The standard is applicable to commercial, scientific, or military software which runs on any digital computer. The standard addresses the documentation of both initial development testing and the testing of subsequent software releases.

The standard may be applied to all phases of testing, from module testing through user acceptance. Since all of the documents may not be useful in each test phase, the particular documents to be used in a phase are not specified. Each organization using the standard will need to specify the classes of software to which the standard applies and the specific documents to be generated for a particular test phase.

2.	Items outside the scope of the standard

The standard does not call for specific testing methodologies, approaches, techniques, facilities, or tools and does not specify the documentation of the use of the items. Additional test documentation may be required. The standard also does not imply or impose specific methodologies for documentation control, configuration management, or quality assurance. Within each standard document, the content of each section may be tailored to the particular application and the particular testing phase. Additional documents may be added to the basic set, additional sections may be added to any document, and additional content may be added to any section. Some or all of the content of  a section may be contained in another document which is then referenced. Each organization using the standard should specify additional content requirements and conventions in order to reflect their own particular methodologies, approaches, facilities, and tools for testing, documentation control, configuration management, and quality assurance.

The standard applies to documentation on electronic media as well as on paper. For those documents requiring approval signatures,, paper must be used unless the electronic documentation system has a secure approval annotation mechanism and that mechanism is used.

3.	Test Plan

3.1	Purpose.

To prescribe the scope, approach, resources, and schedule of the testing activities.

To identify the items being tested, the features to be tested, the testing tasks to be performed, the personnel responsible for each task, and the risks associated with the test plan.

3.2	Outline.

A test plan shall have the following structure:

1.	Test-plan identifier

2.	Introduction

3.	Test items

4.	Features to be tested

5.	Features not to be tested

6.	Approach

7.	Item pass/fail criteria

8.	Suspension criteria and resumption requirements

9.	Test deliverables

10.	Testing tasks

11.	Environmental needs

12.	Responsibilities

13.	Staffing and training needs

14.	Schedule

15.	Risks and contingencies

16.	Approvals

3.2.1	Test-Plan Identifier

3.2.2	Introduction

References to the following documents, when they exist, are required in the highest level test plan:

Project Authorization

Project Plan

Quality Assurance Plan

Configuration Management Plan

Relevant Policies

Relevant Standards

3.2.3	Test Items

Identify the test items including version/revision levels.

Reference the following item documentation, if it exists, for each item:

Requirements Specification

Design Specification

User’s Guide

Operations Guide

Installation Guide

3.2.4	Features to be Tested

3.2.5	Features Not to be Tested

3.2.6	Approach

3.2.7	Item Pass/Fail Criteria

3.2.8	Suspension Criteria and Resumption Requirements

3.2.9	Test Deliverables

The following documents should be included:

Test Plan

Test Design Specifications

Test Case Specifications

Test Procedure Specifications

Test Item Transmittal Reports

Test Logs

Test Incident Reports

Test Summary Reports

Test Input Data and Test Output Date

Test Tools

3.2.10	Test Tasks

3.2.11	Environmental Needs

3.2.12	Responsibilities

3.2.13	Staffing and Training

3.2.14	Schedule

3.2.15	Risks and Contingencies

3.2.16	Approvals

4.	Test-Design Specification

4.1	Purpose

Specify refinements of the test approach

Identify the features to be tested by the design and its associated tests.

4.2	Outline

1.	Test-design-specification identifier

2.	Features to be tested

3.	Approach refinements

4.	Test identification

5.	Feature pass/fail criteria

4.2.1	Test-Design-Specification Identifier

4.2.2	Features to be Tested

Identify test items and describe the features and combination of features which are the object of this design specification.

For each feature or features combination, a reference to its associated requirements in the item requirement specification or design description should be included.

4.2.3	Approach Refinements

Specify refinements to the approach described in the test plan.

Include specific test techniques to be used.

Identify method of analyzing test results.

4.2.4	Test Identification

List the identifier and a brief description of each test case associated with this design.

4.2.5	Feature Pass/Fail Criteria

5.	Test-Case Specification

5.1	Purpose

To define a test case identified by a test design specification

5.2	Outline

A test-case specification shall have the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

5.2.1	Test-Case-Specification Identifier

5.2.2	Test Items

For each test item, reference the following documentation:

Requirements specification

Design specification

Users guide

Operations guide

Installation guide

5.2.3	Input Specifications

Specify each input required to execute the test case, by value, by filename, etc.

Identify all appropriate data bases, files, terminal messages, memory resident areas, etc.

Specify all relationships between inputs.

5.2.4	Output Specifications

Specify all the outputs and features required of the test items.

5.2.5	Environmental Needs

5.2.5.1		Hardware

5.2.5.2		Software

5.2.3		Other environmental requirements

5.2.6	Special Procedural Requirements

5.2.7	Intercase Dependencies

6.	Test-Procedure Specifications

6.1	Purpose

To specify the steps for executing a set of test cases or, more generally, the steps used to analyze a software item in order to evaluate a set of features.

6.2	Outline

A test-procedure specification shall have the following structure:

1.	Test-procedure-specification identifier

2.	Purpose

3.	Special requirements

4.	Procedure steps

6.2.1.	Test-procedure-specification identifier

Specify the unique identifier assigned to this test-procedure specification. Reference the associated test-design specification.

6.2.2	Purpose

Describe the purpose of the procedure. If the procedure executes any test cases, reference each test-case specification to be executed.

6.2.3	Special Requirements

Identify any special requirements that are necessary to execute the procedure. Identify any prerequisite procedures, special skill requirements, and special environment requirements.

6.2.4	Procedure Steps

6.2.4.1		Log

Describe  any specific methods or formats for logging the results of the test executions, incidents observed, and any other events pertinent to the test.

6.2.4.2		Set Up

Describe the actions necessary to prepare for executing the procedure.

6.2.4.3		Start

Describe the actions necessary to start execution of the test procedure.

6.2.4.4		Proceed

Describe any actions necessary during execution of the procedure.

6.2.4.5		Measure

Describe how any quantitative measurements will be made.

6.2.4.6		Shut Down

Describe the actions necessary to suspend testing, when unscheduled events dictate.

6.2.4.7		Restart

Identify any procedural restart points and describe the actions necessary to restart the procedure at each of these points.

6.2.4.8		Stop

Describe the actions necessary to bring the test procedure execution to an orderly halt.

6.2.4.9		Wrap Up

Describe the actions necessary to restore the environment to its pre-test condition.

6.2.4.10	Contingencies

Describe the actions necessary to deal with anomalous events that may occur during the test execution.

7.	Test-Item Transmittal Report

7.1	Purpose

To identify the test items being transmitted for testing. Identifies the person responsible for each item, its physical location, and its status. Variations from the current item requirements and design are noted in this report.

7.2	Outline

A test-item transmittal report shall have the following structure:

(1)	Transmittal report identifier

(2)	Transmitted items

(3)	Location

(4)	Status

(5)	Approvals

7.2.1	Transmittal-Report Identifier

Specify the unique identifier assigned to this test-item transmittal report.

7.2.2	Transmitted Items

Identify the test items being transmitted, including their version/revision level. Supply references to the item documentation and the test plan relating to the transmitted items. Indicate the people responsible for the transmitted items.

7.2.3	Location

Identify the location of the transmitted item. Identify the media containing the transmitted items, including appropriate labeling or identification.

7.2.4	Status

Describe the status of the test items being transmitted. Include deviations from previous transmittals of these items, and from the test plan. List the incident reports that are expected to be resolved by the transmitted items. Indicate if there are pending modifications to item documentation which may affect the items listed in this transmittal report.

7.2.5	Approvals

Specify the names and titles of all persons who must approve the transmittal.

8.	Test Log

8.1	Purpose

To provide a chronological record of relevant details about the execution of tests.

8.2	Outline

A test log shall have the following structure:

(1)	Test log identifier

(2)	Description

(3)	Activity and event entries

8.2.1	Test-Log Identifier

Specify the unique identifier assigned to this test log.

8.2.2	Description

Information which applies to all entries in the log, such as:

(1) Identify the items being tested including version/revision levels. For each item, supply a reference to its transmittal report.

(2) Identify the attributes of the environments in which the testing is conducted, Include facility identification, hardware being used, system software used, and system resources available.

8.2.3	Activity and Event Entries

For each event, including the beginning and end of activities, record the occurrence date and time and the identity of the author. Include:

8.2.3.1	Execution Description

Record the identifier of the test procedure being executed and supply a reference to the test specification. Record all responsible personnel present during the execution and their functions.

8.2.3.2	Procedure Results

For each execution, record the visually observable results (for example, error messages generated, aborts, and requests for operator action). Also record the location of any output.

8.2.3.3	Environmental Information

Record any environmental conditions specific to this entry.

8.2.3.4	Anomalous Events

Record what happened before and after an unexpected event occurred. Record circumstances surrounding the inability to begin execution of a test procedure or failure to complete a test procedure.

8.2.3.5	Incident-Report Identifiers

Record the identifier of each test-incident report, whenever one is generated.

9.	Test Incident Report

9.1	Purpose

To document any event that occurs during the testing process which requires investigation.

9.2	Outline

A test-incident report shall have the following structure

(1)	Test-incident report identifier

(2)	Summary

(3)	Incident description

(4)	Impact

9.2.1	Test-Incident Report Identifier

Specify the unique identifier assigned to this test incident report.

9.2.2	Summary

Summarize the incident. Identify the test items involved indicating their version/revision level. Reference the appropriate test-procedure specification, test-case specification, and test log.

9.2.3	Incident Description

Provide a description of the incident. The incident description should include the following items:

Inputs

Expected results

Anomalies

Date and time

Procedure step

Environments

Attempts to repeat

Testers

Observers

Related activities and observations that may help to isolate and correct the cause of the incident should be included.

9.2.4	Impact

If known, indicate the impact that this incident will have on test plans, test-design specifications, test-procedure specifications, or test-case specifications.

10.	Test-Summary Report

10.1	Purpose

To summarize the results of the designated activities and to provide evaluations based on these results.

10.2	Outline

A test-summary report shall have the following structure:

(1)	Test-summary report identifier

(2)	Summary

(3)	Variances

(4)	Comprehensive assessment

(5)	Summary of results

(6)	Evaluation

(7)	Summary of activities

(8)	Approvals

10.2.1	Test-Summary Report

Specify the unique identifier assigned to this test-summary report.

10.2.2	Summary

Summarize the evaluation of the test items. Identify the items tested, indicating version.revision level. Indicate the environment in which the testing activities took place. For each test item, reference the following documents if they exist: test plan, test-design specification, test-procedure specifications, test-item transmittal reports, test logs, test-incident reports.

10.2.3	Variances

Report any variances of the test items from their design specifications. Indicate any variances from the test plan, test designs, or test procedures.

10.2.4	Comprehensiveness Assessment

Evaluate the comprehensiveness of the testing process against the comprehensiveness criteria specified in the test plan. Identify features or feature combinations which were not sufficiently tested and explain the reasons.

10.2.5	Summary of results

Summarize the results of testing. Identify all resolved incidents and summarize their resolutions. Identify all unresolved incidents.

10.2.6	Evaluation

Provide an overall evaluation of each test item including its limitations. This evaluation must be based upon the test results and the item level pass/fail criteria.

10.2.7	Summary of Activities

Summarize the major testing activities and events. Summarize resource consumption data, for example, total staffing level, total machine time, and total elapsed time used for each of the major testing activities.

10.2.8	Approvals

Specify the names and titles of all persons who must approve this report.