Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Abort Transfer

Document ID:		1.2.840.113702.1.3.7.2

Version Date:		1/29/96

Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Abort Transfer

i)	Purpose:

This Test-Case Specification document defines a test of the behavior of the Picker PQ2000 CT scanner (CT Storage SCU) when an association is aborted locally.

ii)	Outline:

This test-case specification has the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1. Test-Case-Specification Identifier

This example Test Case Specification has the ID:

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| Abort Transfer | Local abort of a transfer | 1.2.840.113702.1.3.7.2 |

2. Test Items

The behavior of the HANI Application Entity as implemented in Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner will be tested under the condition of a transfer abort issued from the host user interface.

3. Input Specifications

At the PQ2000 operatorÕs console select HyperLan transfer destination, select a study from the directory and initiate transfer. When the transfer is approximately 1/2 complete, issue the abort command.

4. Output Specifications

Monitor and record the status display which must indicate transfer aborted. Observe that control is returned normally.

5. Environmental Needs

5.1	Hardware

A Picker PQ2000 CT scanner operatorÕs console networked together with a system that acts as a CT storage SCP.

5.2	Software

Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner

5.3	Other environmental requirements

A telephone must be accessible to the operator of the CT Scanner console and of the SCP so that transfer start events can be communicated

6. Special Procedural Requirements

Two trained test personnel are required.

7. Intercase Dependencies

If the association is not cleanly and completely aborted, there may be an impact on subsequent association establishment.