Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level_3 Test-Case Specification Example

SCP Disk Full Resource Boundary

Document ID:		1.2.840.113702.1.3.11.2

Version Date:		1/29/96

Application Entities Interconnectivity

Level_3 Test-Case Specification Example

SCP Disk Full Resource Boundary

i)	Purpose:

This Test-Case Specification document defines a test of boundary condition behavior in the SCP Disk Full case, for the Picker PQ2000 CT scanner (CT Storage SCU).

ii)	Outline:

This test-case specification has the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1.	Test-Case-Specification Identifier

This example Test Case Specification has the ID

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| SCP Disk Full Boundary | CT Image Storage Services attempted transfer with SCP Disk Full | 1.2.840.113702.1.3.11.2 |

2.	Test Items

The ability of the HANI Application Entity as implemented in Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner to correctly detect and report an SCP Disk Full error state will be tested.

3.	Input Specifications

Select a CT exam with approximately 20 images.

At the PQ2000 operator's console select HyperLan transfer destination and initiate transfer.

4.	Output Specifications

Monitor and record the status display which must indicate transfer failure do to SCP resource limitation.

5.	Environmental Needs

5.1		Hardware

A Picker PQ2000 CT scanner operator's console networked together with a system that acts as a CT storage SCP. The SCP must have insufficient storage to accept all 20 images being transferred.

5.2		Software

Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner.

5.3		Other environmental requirements

A telephone must be accessible to the operator of the CT Scanner console and of the SCP so that transfer start events can be communicated

6.	Special Procedural Requirements

Two trained test personnel are required.

7.	Intercase Dependencies

None.