Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Repeat Transfer

Document ID:		1.2.840.113702.1.3.6.2

Version Date:		1/29/96

Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Repeat Transfer

i)	Purpose:

This Test-Case Specification document defines a test of the behavior of the Picker PQ2000 CT scanner (CT Storage SCU) when a successful transfer is repeated.

ii)	Outline:

This test-case specification has the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1. Test-Case-Specification Identifier

This example Test Case Specification has the ID:

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| Repeat Transfer | Transfer a study twice | 1.2.840.113702.1.3.6.2 |

2. Test Items

The behavior of the HANI Application Entity as implemented in Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner will be tested under the condition of repeated transfer of a set of images.

3. Input Specifications

At the PQ2000 operatorÕs console select HyperLan transfer destination, select study m for  Patient n from the directory and initiate transfer. When the transfer completes, repeat using the same study.

4. Output Specifications

Monitor and record the status display which must indicate transfer complete for the first transfer and a error status (warning) for the second.

5. Environmental Needs

5.1	Hardware

A Picker PQ2000 CT scanner operatorÕs console networked together with a system that acts as a CT storage SCP.

5.2	Software

Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner

5.3	Other environmental requirements

A telephone must be accessible to the operator of the CT Scanner console and of the SCP so that transfer start events can be communicated

6. Special Procedural Requirements

Two trained test personnel are required.

7. Intercase Dependencies

None.