Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Association Establishment and Image Transfer, Single Study

Document ID:		1.2.840.113702.1.3.15.1

Version Date:		1/29/96

Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Association Establishment and Image Transfer, Single Study

i)	Purpose:

This Test-Case Specification document defines a test of basic DICOM Association Establishment and CT image transfer for the General Electric Advantage Review workstation (CT Storage SCP).

ii)	Outline:

This test-case specification has the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1.	Test-Case-Specification Identifier

This example Test Case Specification has the ID

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| Association Establishment and Image Transfer | Association establishment and CT Image Storage Services for a Single Study | 1.2.840.113702.1.3.15.1 |

2.	Test Items

Basic Association Establishment and CT image storage functions of the Local DICOM Import Application Entity as implemented in Software Revision 1.03 of the General Electric Advantage Review Workstation will be tested.

3.	Input Specifications

The user interaction input step at the Service Class User Application Entity is that required to cause the Application Entity to issue a DIMSE C-STORE  request for the Service Object Pair under test.

4.	Output Specifications

Monitor and record the status display and observe the patient directory upon completion of transfer. The status display must indicate transfer complete and the study should appear in the directory.

5.	Environmental Needs

5.1		Hardware

A GE Advantage Review Workstation networked together with a system that acts as a CT storage SCU.

5.2		Software

Advantage Review Software Revision 1.03.

5.3		Other environmental requirements

A telephone must be accessible to the operator of the Advantage Review workstation and of the SCU so that transfer start events can be communicated.

6.	Special Procedural Requirements

Two trained test personnel are required. The GE workstation status display updates rapidly overwriting previous messages. It must be monitored very carefully.

7.	Intercase Dependencies

The available disk storage on the system must be monitored after each transfer to avoid resource limit errors.