Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Association and CT Storage Services

Level_3 Test-Design Specification Example

Document ID:		1.2.840.113702.1.3.2.1

Version Date:		111/12/95

Application Entities Interonnectivity

Association and CT Storage Services

Level_3 Test-Design Specification Example

i)	Purpose

The purpose of this Test Design Specification document is to detail refinements of the test approach presented in the Test Plan for interconnectivity testing between the Picker PQ2000 CT scanner and the General Electric Advantage Review Workstation. The Test Design Specification  will identify the features to be tested by the design and its associated tests, identify procedures for executing the defined tests, and identify specific test cases.

This document is the Level_3 Test Design example document for the combined Association Services and CT Storage Service Class. Association Services and Storage are operationally inseparable for the Picker CT SCU and therefore the Test Plan, Test Design and Test Case documents have been written to reflect this operational reality. Interconnectivity testing between the Picker CT and GE workstation does not require multiple separable Test Design documents, nor is it necessary to have a separable Test Procedure document. This document is intended to represent a specific Test Design and Procedure which accommodates all of the characteristics of the two specified, interconnected Application Entities both of which comply with the DICOM Standard, ACR/NEMA Standards Publication No. PS 3.

ii)	Outline

A test-design specification document shall have the following structure.

1.	Test-design-specification identifier

2.	Features to be tested

3.	Approach refinements

4.	Test identification

5.	Feature pass/fail criteria

6.	Special Requirements

7. 	Procedure Steps

1.	Test-Design-Specification Identifier

The Test-Design-Specification Identifier shall be a unique identifier for both suppliers of the Application Entities to be tested. This example Test Design Specification has the ID: 1.2.840.113702.1.3.2.1

2.	Features to be Tested

The features which are testable and the design of tests for those features is completely dependent on the access that the testers have to enter and observe information in the systems under test. In this case, the systems under test were a Picker PQ2000 CT Scanner and a General Electric Advantage Review Workstation. The Picker CT scanner is strictly a Service Class User of the CT Store SOP. While the GE workstation is capable of being both a Service Class Provider and a Service Class User of the CT Store SOP, this cross-vendor testing situation allowed for testing only the Service Class Provider SOP capability of this system. The operational granularity and the access to logging information that is accessible to testers of these two systems differs considerably from the idealized model presented in the Testing Guidelines.

The following features will be tested between the two Application Entities (the HANI Application Entity as implemented in Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner and the Local DICOM Import Application Entity as implemented in Software Revision 1.03 of the General Electric Advantage Review Workstation):

Basic TCP/IP connectivity

Basic Association Establishment and image transfer

SCU role of the CT Image Storage SOP Class (Picker)

SCP role of the CT Image Storage SOP Class (General Electric)

Association Establishment Policies (Picker):

Single Association Mode

Multiple Association Mode

Association Abort

Behavior at Boundary Cases

Repeatability

The Picker PQ2000 CT scanner operates in two modes: Single Association and Multiple Association. In Single Association Mode only one Association may be established at a time. When the Study Identifier changes in the image stream, the HANI Application Entity releases the existing Association and established a new one. In Multiple Association Mode, HANI creates a new Association when the Study Identifier changes and proceeds to transfer images over multiple Associations in parallel.

3.	Approach Refinements

In the Picker CT scanner, the most primitive operation available for testing is the one that operators would clearly prefer. That operation allows the operator to select one or more studies, (a set of images), and then to ‘Send’ the images to the Service Class Provider AE. In principle, the Picker SCU will attempt to establish multiple concurrent associations if several studies are requested simultaneously. The feedback available to the operator is succinct. The Picker SCU counts the number of images transferred and provides an operator with the information of the number of images transferred and whether the entire operation was successful or failed at some point.

While the Picker SCU does provide output messages to the operator, no additional log files were detected. The messages visible to an operator are:

1.	Transferring image .....

2(a).	HyperLAN Network Transfer Complete

2(b).	Error in Network Call, Quit Now.

The GE Workstation provided a bit more detailed information that was useful for testing, but that information is presented in a window printout that is continually updated and partially overwritten and is apparently not logged.  The difficulty, therefore, is to catch the messages going by, not miss any, and to correctly interpret the sequence. From the associated Conformance Statement, and from all observations, the GE Workstation does not establish concurrent associations. Each message displayed to the operator is date and time stamped, but for simplicity the date/time information will not be included in sample messages described herein. As images are received in DICOM format, the GE Workstation takes each message in sequence and converts the image data to an internal format and stored in the local database. That conversion does proceed concurrent with receiving additional images, and messages indicating each conversion are also displayed to the window. Those messages will also be omitted from our testing description since they are outside the scope of DICOM interconnectivity testing. It should be noted, however, that the information output of the conversion process is extremely relevant to DICOM Interoperability, as defined in the Introduction and Overview Document of the Testing Guidelines.

The testing relevant messages displayed by the GE Workstation are:

1. Connection (date time) Starting

2. Connection (date time) Running DICOM:  Receiving Images

3. Connection (date time) Running  DICOM Import Received :  n Images from Application Entity Name.

4. Connection (date time) Exiting  DICOM Import Received :  n Images from Application Entity Name.

5. Connection (date time) Abort DICOM Import Received :  n Images from Application Entity Name.

4.	Test Identification

By analysis of the conformance statements of the Picker PQ2000 CT scanner and General Electric Advantage Review Workstation the following interconnectivity tests have been designed:

| Feature | Test Case ID | Test Case Description |
| --- | --- | --- |
| Basic TCP/IP connectivity | 1.2.840.113702.1.3.3.1 | Validation of TCP configuration and connectivity |
| Basic Association Establishment and image transfer (SCU and SCP roles) | 1.2.840.113702.1.3.4.1 | Single Association Mode, single study |
| Association Establishment - Single Association Mode | 1.2.840.113702.1.3.5.1 | Select/Queue multiple cases; 3 different patients, 4 studies ( 2 studies of one patient) |
| Repeatability of Transfers | 1.2.840.113702.1.3.6.1 | Repeatability Test; Single patient, resend the study |
| Association Abort | 1.2.840.113702.1.3.7.1 | Abort; Start sending a long series of images and Abort |
| Association Establishment - Multiple Association Mode | 1.2.840.113702.1.3.8.1 | Multiple Association Mode, single study |
| Association Establishment - Multiple Association Mode | 1.2.840.113702.1.3.9.1 | Multiple Association Mode, multiple study |
| Boundary Behavior - Multiple Simultaneous Associations | 1.2.840.113702.1.3.10.1 | Queue up multiple cases and try to force simultaneous associations |
| Boundary Behavior - Resource boundary | 1.2.840.113702.1.3.11.1 | Full disk |

5.	Feature Pass/Fail Criteria

| Feature | Test Case ID | Pass Criteria |
| --- | --- | --- |
| Basic TCP/IP connectivity | 1.2.840.113702.1.3.3.1 | node is alive |
| Basic Association Establishment and image transfer (SCU and SCP roles) | 1.2.840.113702.1.3.4.1 | Picker: HyperLAN Transfer Complete GE: DICOM Import Received : n Images from Application Entity Name |
| Association Establishment - Single Association Mode | 1.2.840.113702.1.3.5.1 | For each case - Picker: HyperLAN Transfer Complete GE: DICOM Import Received : n Images from Application Entity Name |
| Repeatability of Transfers | 1.2.840.113702.1.3.6.1 | Picker: HyperLAN Transfer Complete GE: DICOM Import Received : n Images from Application Entity Name or warning message indicating previously stored UID. |
| Association Abort | 1.2.840.113702.1.3.7.1 | Picker: HyperLAN Transfer Complete GE: DICOM transfer aborted error message |
| Association Establishment - Multiple Association Mode | 1.2.840.113702.1.3.8.1 | Picker: HyperLAN Transfer Complete GE: DICOM Import Received : n Images from Application Entity Name |
| Association Establishment - Multiple Association Mode | 1.2.840.113702.1.3.9.1 | For each case - Picker: HyperLAN Transfer Complete GE: DICOM Import Received : n Images from Application Entity Name |
| Boundary Behavior - Multiple Simultaneous Associations | 1.2.840.113702.1.3.10.1 | For each case - Picker: HyperLAN Transfer Complete GE: DICOM Import Received : n Images from Application Entity Name |
| Boundary Behavior - Resource boundary | 1.2.840.113702.1.3.11.1 | Picker: Error in Network Call, Quit Now. GE: disk full error condition |

6.	Special Requirements

The Picker CT scanner and General Electric workstation were physically connected to a common network. Both TCP/IP and DICOM configuration was completed such that it was possible to ÒpingÓ each machine from the other.  Domain and name services were provided on the network so that a host name to IP address translation was supported for both systems under test.

7. Procedure Steps

7.1 Log

Create a test log data sheet and record all system states responses from each system under test.

7.2 Setup

Configure TCP/IP communication between the systems under test. Add appropriate Application Entity 	Titles, IP addresses and port numbers to the database of each system. Configure GE workstation in 	configuration 1, no simultaneous associations.

7.3 Measure

Configure the Picker CT scanner to allow only one association to be established at a time.

1.	Single study

2.	Select/Queue multiple cases; 3 different patients, 4 studies ( 2 studies of one patient)

3.	Repeatability Test;  Single patient, resend the study

4.	Abort; Start sending a long series of images and Abort

Configure the Picker CT scanner to allow simultaneous associations.

5.	Single study

6.	Select/Queue multiple cases; 3 different patients, 4 studies ( 2 studies of one patient)

7.	Queue up multiple cases and try to force simultaneous associations.

7.4 Wrap up

Compile Test Summary.  Submit Test Summary and Test-Design Specification for Web publication.

7.5 Contingencies

Compile and submit Test Incident Reports to vendors. If remedies or work-arounds are possible, implement and retest.