Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Storage Service Class

Level_2 Test-Design Specification Template

Document ID:		1.2.840.113702.1.2.4.1

Version Date:		10/20/95

Application Entities Interconnectivity

Storage Service Class

Level_2 Test-Design Specification Template

i)	Purpose

The purpose of the Test Design Specification document is to detail refinements of the test approach presented in the Test Plan.

To identify the features to be tested by the design and its associated tests.

ii)	Outline

A test-design specification document shall have the following structure.

1.	Test-design-specification identifier

2.	Features to be tested

3.	Approach refinements

4.	Test identification

5.	Feature pass/fail criteria

1.	Test-Design-Specification Identifier

The Test-Design-Specification Identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

2.	Features to be Tested

Identify test items and describe the features and combination of features which are the object of this design specification.

For each feature or features combination, a reference to its associated requirements in the item requirement specification or design description should be included.

In testing the interconnectivity of the Storage Service Class between two Application Entities, the DIMSE C-STORE service will be tested for all Service-Object Pairs (SOPs) that are mutually supported by both Application Entities.

The following features will be tested in every test of interconnectivity of the Association Services between two Application Entities.

Storage Service:

C-STORE	Store a composite SOP Instance from the Service Class User (SCU) to the Service Class Provider(SCP).

Standard SOP Classes:

Computed Radiography Image Storage

CT Image Storage

Ultrasound Multi-frame Image Storage

MR Image Storage

Nuclear Medicine Image Storage

Ultrasound Image Storage

Secondary Capture Image Storage

Stand-alone Curve Storage

Stand-alone Modality LUT Storage

Stand-alone VOI LUT Storage

The documentation associated with each test should include:

Test Design Specification Number

Identification of the Conformanace Statements with specific reference to the relevant sections in the Conformance Statements most concerned with the specific test.

Identification of other relevant documentation of the Application Entities with specific reference to the relevant sections of the documentation most concerned with the specific test.

3.	Approach Refinements

Specify refinements to the approach described in the test plan.

Include specific test techniques to be used.

Identify method of analyzing test results.

The approach used for testing the Storage Service Class involves repeated exercising of the DIMSE C-STORE service for each of the SOPs mutually supported by both of the Application Entities. The assumptions are:

1.	The  C-STORE service will be invoked from the user interface level of operation

2.	Each invocation of the DIMSE service primitives is logged

3.	Information about the completion ofa C-STORE service, successful or unsucessful, may or may not be passed to the user interface level of operation

The actual design of tests will depend on the the amount of information and the control that the designers and implementers have made available to the system operators. In testing for interconnectivity of the Storage Service Class, there are several levels of information access that may be available to the user:

Event Access:	Access to the information that a significant event has occurred.

Command Data Access:	Access to some or all of the command data

Data Set Access:	Access to the information contained in the data set of the messages.

The assumption to be made in this template version of the test design is that at least Event Access and Command Data Access levels of information are available for testing. The approach taken for each test will be to run the test and to ensure first that the significant primitive events have taken place in the proper sequence. Then the Command Data level can be examined and, finally,  the Data Access level of information, if available, can be examined.

The primary method of analyzing Event Access test results will be by examining ‘log’ files, both at the application level of invoking the DIMSE C-STORE service and at the level of the DIMSE service providers.

4.	Test Identification

List the identifier and a brief description of each test case associated with this design.

Each Test Identifier shall be a unique identifier for both providers of the Application Entities to be tested.

At this level of testing, the C-STORE service will be tested for all of the Service Object Pairs (SOPs) the are mutually supported by the Application Entities for the Storage Service Class. A test identification shall be established for each Service-Object Pair to be tested between the two Application Entities.

5.	Feature Pass/Fail Criteria

The Pass/Fail criterion for Storage Service Class interconnectivity of two Application Entities is that they should exhibit the behavior expressed in the DICOM Standard, PS 3, Part 4 and Part 7. For each invocation of the storage service, the primitive operations should take place in the proper sequence and each of the Application Entities should be left in the appropriate state relative to the association. The Commamd Set data, if available, can be chacked to confirm that both Application Entities indicate the same outcome. Ultimately, the test of interconnectivity is clinical utility of the Data Set transferred for storage.