Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Association and CT Storage Services

Level_3 Test-Design Specification Example

Document ID:		1.2.840.113702.1.3.2.2

Version Date:		1/20/96

Application Entities Interconnectivity

Association and CT Storage Services

Level_3 Test-Design Specification Example

i)	Purpose

The purpose of this Test Design Specification document is to detail refinements of the test approach presented in the Test Plan for interconnectivity testing of the Picker PQ 2000 CT scanner. The Test Design Specification  will identify the features to be tested by the design and its associated tests, identify procedures for executing the defined tests, and identify specific test cases.

This document is the Level_3 Test Design example document for the combined Association Services and CT Storage Service Class. Association Services and Storage are operationally inseparable for the Picker CT SCU and therefore the Test Plan, Test Design and Test Case documents have been written to reflect this operational reality. DICOM interconnectivity testing of the Picker CT scanner does not require multiple separable Test Design documents, nor is it necessary to have a separable Test Procedure document. This document is intended to represent a specific Test Design and Procedure which accommodates all of the characteristics of the specified Application Entity, which complies with the DICOM Standard, ACR/NEMA Standards Publication No. PS 3.

ii)	Outline

A test-design specification document shall have the following structure.

1.	Test-design-specification identifier

2.	Features to be tested

3.	Approach refinements

4.	Test identification

5.	Feature pass/fail criteria

6.	Special Requirements

7. 	Procedure Steps

1.	Test-Design-Specification Identifier

This example Test Design Specification has the ID: 1.2.840.113702.1.3.2.2

2.	Features to be Tested

The features which are testable and the design of tests for those features is completely dependent on the access that the testers have to enter and observe information in the systems under test. The Picker CT scanner is strictly a Service Class User of the CT Store SOP. The operational granularity and the access to logging information that is accessible to testers differs considerably from the idealized model presented in the Testing Guidelines.

The following features of the HANI Application Entity as implemented in Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner will be tested:

Basic TCP/IP connectivity

Basic Association Establishment and image transfer

SCU role of the CT Image Storage SOP Class

Association Establishment Policies:

Single Association Mode

Multiple Association Mode

Association Abort

Behavior at Boundary Cases

Repeatability

The Picker PQ 2000 CT scanner operates in two modes: Single Association and Multiple Association. In Single Association Mode only one Association may be established at a time. When the Study Identifier changes in the image stream, the HANI Application Entity releases the existing Association and established a new one. In Multiple Association Mode, HANI creates a new Association when the Study Identifier changes and proceeds to transfer images over multiple Associations in parallel.

3.	Approach Refinements

In the Picker CT scanner, the most primitive operation available for testing allows the operator to select one or more studies, (a set of images), and then to ‘Send’ the images to the Service Class Provider AE. In principle, the Picker SCU will attempt to establish multiple concurrent associations if several studies are requested simultaneously. The feedback available to the operator is succinct. The Picker SCU counts the number of images transferred and provides an operator with the information of the number of images transferred and whether the entire operation was successful or failed at some point.

While the Picker SCU does provide output messages to the operator, no additional log files were detected. The messages visible to an operator are:

1.	Transferring image .....

2(a).	HyperLAN Network Transfer Complete

2(b).	Error in Network Call, Quit Now.

4.	Test Identification

By analysis of the conformance statement of the Picker PQ 2000 CT scanner the following interconnectivity tests have been designed:

| Feature | Test Case ID | Test Case Description |
| --- | --- | --- |
| Basic TCP/IP connectivity | 1.2.840.113702.1.3.3.2 | Validation of TCP configuration and connectivity |
| Basic Association Establishment and image transfer (SCU role) | 1.2.840.113702.1.3.4.2 | Single Association Mode, single study |
| Association Establishment - Single Association Mode | 1.2.840.113702.1.3.5.2 | Select/Queue multiple cases; 3 different patients, 4 studies ( 2 studies of one patient) |
| Repeatability of Transfers | 1.2.840.113702.1.3.6.2 | Repeatability Test; Single patient, resend the study |
| Association Abort | 1.2.840.113702.1.3.7.2 | Abort; Start sending a long series of images and Abort |
| Association Establishment - Multiple Association Mode | 1.2.840.113702.1.3.8.2 | Multiple Association Mode, single study |
| Association Establishment - Multiple Association Mode | 1.2.840.113702.1.3.9.2 | Multiple Association Mode, multiple study |
| Boundary Behavior - Multiple Simultaneous Associations | 1.2.840.113702.1.3.10.2 | Queue up multiple cases and try to force simultaneous associations |
| Boundary Behavior - Resource boundary | 1.2.840.113702.1.3.11.2 | Full disk |

5.	Feature Pass/Fail Criteria

| Feature | Test Case ID | Pass Criteria |
| --- | --- | --- |
| Basic TCP/IP connectivity | 1.2.840.113702.1.3.3.2 | node is alive |
| Basic Association Establishment and image transfer (SCU role) | 1.2.840.113702.1.3.4.2 | HyperLAN Transfer Complete |
| Association Establishment - Single Association Mode | 1.2.840.113702.1.3.5.2 | For each case - HyperLAN Transfer Complete |
| Repeatability of Transfers | 1.2.840.113702.1.3.6.2 | HyperLAN Transfer Complete |
| Association Abort | 1.2.840.113702.1.3.7.2 | HyperLAN Transfer Complete |
| Association Establishment - Multiple Association Mode | 1.2.840.113702.1.3.8.2 | HyperLAN Transfer Complete |
| Association Establishment - Multiple Association Mode | 1.2.840.113702.1.3.9.2 | For each case - HyperLAN Transfer Complete |
| Boundary Behavior - Multiple Simultaneous Associations | 1.2.840.113702.1.3.10.2 | For each case - HyperLAN Transfer Complete |
| Boundary Behavior - Resource boundary | 1.2.840.113702.1.3.11.2 | Error in Network Call, Quit Now. |

6.	Special Requirements

The Picker CT scanner must be physically connected to a local area network. Both TCP/IP and DICOM configuration must be established such that it is possible to ÒpingÓ the Picker machine from another network device.  Domain and name services must be provided on the network so that a host name to IP address translation is supported for all systems under test.

7. Procedure Steps

7.1 Log Create a test log data sheet and record all system status responses.

7.2 Setup Configure TCP/IP communication. Add appropriate Application Entity Titles, IP addresses and 	port numbers to the database.

7.3 Measure Configure the Picker CT scanner to allow only one association to be established at a time.

1.	Single study

2.	Select/Queue multiple cases; 3 different patients, 4 studies ( 2 studies of one patient)

3.	Repeatability Test;  Single patient, resend the study

4.	Abort; Start sending a long series of images and Abort

Configure the Picker CT scanner to allow simultaneous associations.

5.	Single study

6.	Select/Queue multiple cases; 3 different patients, 4 studies ( 2 studies of one patient)

7.	Queue up multiple cases and try to force simultaneous associations.

7.4 Wrap up Compile Test Summary.

7.5 Contingencies Compile and submit Test Incident Reports. If remedies or work-arounds are possible, implement and retest.