Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level 1 Test Plan Template

Master Test Plan

Document ID:		1.2.840.113702.1.1.2.1

Version Date:		10/20/95

Application Entities Interconnectivity

Level_1 Test Plan Template

Master Test Plan

i.	Purpose.

The purpose of a Test Plan document is to prescribe the scope, approach, resources, and schedule of the testing activities.

To identify the items being tested, the features to be tested, the testing tasks to be performed, the personnel responsible for each task, and the risks associated with the test plan.

This document is the Level_1 Test Plan template document to address complete testing for interconnectivity between two Application Entities and is intended to be the template base document from which other test plans to define such testing are derived. This document is intended to represent a generic Test Plan which accommodates all of the characteristics of interconnected Application Entities that follow the DICOM Standard, ACR/NEMA Standards Publication No. PS 3.

ii.	Outline.

A test plan document shall have the following structure:

1.	Test-plan identifier

2.	Introduction

3.	Test items

4.	Features to be tested

5.	Features not to be tested

6.	Approach

7.	Item pass/fail criteria

8.	Suspension criteria and resumption requirements

9.	Test deliverables

10.	Testing tasks

11.	Environmental needs

12.	Responsibilities

13.	Staffing and training needs

14.	Schedule

15.	Approvals

1.	Test-Plan Identifier

The Test Plan Identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

2.	Introduction

2.1	Objectives.  A test plan for testing the interconnectivity between two Application Entities shall support the following objectives.

(1)  To detail the activities required to prepare for and support the test.

(2)  To communicate to all responsible parties the tasks which they are to perform and the schedule to be followed in performing the tasks.

(3)  To define the sources of information used to prepare the test plan.

(4)  To define the test tools and environment needed to conduct the test.

2.2	Background.  The DICOM Standard, ACR/NEMA Standards Publication No. PS 3 has been developed to define a standard method of transferring images and associated information between devices manufactured by different vendors. The standard defines various information objects and the operations or services that may be performed on the objects, organizes the objects and services into Service-Object Pair (SOP) Classes, and groups the SOP Classes into Service Classes which can be supported by any Application Entity. The DICOM Standard also allows each implementor of an Application Entity to determine the degree to which the Application Entity will conform to the DICOM Standard for each Service-Object Pair Class. The standard further distinguishes the services that control the ability to exchange information as the Association Services.

2.3	Scope.  A test for interconnectivity between two Application Entities has the objective to show that the Application Entities are able to exchange messages in accordance with the DICOM Standard, ACR/NEMA Standards Publication No. PS 3. This Master Test Plan presents a schema for organizing the testing of two Application Entities for interconnectivity. The plan partitions the testing process in exactly the same way that the DICOM Standard organizes the Service Classes and the Association Services. For each of these, a complete testing cycle will be developed, with the accompanying set of documents required to record the design and execution of the testing process. The Association Services are distinguished for testing because they are required to be operational before any of the Services Classes can be tested.

This Master Test Plan and any other testing documents at this level serve only to present the schema for organizing testing of the Service Classes and the Association Services. The testing documentation at this level does not attempt to present detailed testing information for each of the Association Services or any of the Services Classes. That degree of detail is left to the testing documentation associated with each of the Service Classes and the Association Services. The test plan for each category will include functionality testing of the Application Entities, the operator interfaces for invoking message transfer, and indications of successful and failed operations.

2.4	References.  References to the following documents, when they exist, are required in the test plan:

Project Authorization

Project Plan

Quality Assurance Plan

Configuration Management Plan

Relevant Policies

Relevant Standards

In multilevel test plans, each lower-level plan must reference the next higher-level plan. Specifically, each of the test plans developed for the Association Services and the Service Classes must reference this Master Test Plan.

The following documents shall be included in the references of every test for interconnectivity of two Application Entities:

DICOM Standard, ACR/NEMA Standards Publication No. PS 3

The Conformance Statements of each Application Entities to be tested.

3.	Test Items

The items to be tested for interconnectivity will be the two Application Entities of interest. These test items shall be identified down to the level of version/revision number. As specified by the DICOM Standard, the Upper Layer Association Services are included in the Application Entities and are therefore included in the testing process.

The following documents, if available, will be used to provide the definition of correct operation.

Requirements Specification

Design Specification

User’s Guide

Operations Guide

Installation Guide

4.	Features to be Tested

The following features will be tested in every test of interconnectivity between two Application Entities. For any two Application Entities, only the services which are common to both Application Entities can of course be tested.

Association Services

Supported Service Classes and Service Class Roles

Storage Service Class

Query / Retrieve Service Class

Print Service Class

Patient, Study, and Results Management Service Classes

5.	Features Not to be Tested

6.	Approach

The test personnel will use the DICOM Standard, PS 3, the respective Conformance Statements and documentation of the Application Entities to prepare all test design, test case, and test procedure specifications. The approach in this Master Test Plan will be to organize the testing schema for each of the Service Classes and the Association Services. Detailed expansion of testing details will be developed within the individual categories.

The overall test plan and the design of tests will depend on the amount of information and the control that the designers and implementors have made available to the system operators. In testing for interconnectivity of Application Entities, there are three levels of information access that may be available to the user:

Event Access:	Access to the information that a significant event has occurred.

Command Data Access:	Access to the relevant portions of the command set information.

Data Set Access:	Access to the information contained in the data set of the messages.

The assumption to be made in the template versions of the test plans is that at least Event Access level of information is available for testing. The approach taken for each test will be to run the test and to ensure that the significant events have taken place in the proper sequence. If the Data Access level of information is available, then the data can be checked more closely.

7.	Item Pass/Fail Criteria

The Pass/Fail criterion for interconnection of two Application Entities in each of the Service Classes and the Association Services is that they should exhibit the behavior expressed in the DICOM Standard. For each test in each category, the operational events should take place in the proper sequence and each of the Application Entities should be left in the appropriate state following the operation.

8.	Suspension Criteria and Resumption Requirements

8.1 Suspension Criteria. The Association Services are a prerequisite to the operation of all of the Services Classes and must operate successfully in order for testing of any of the Service Classes to proceed. Should any of the Association Services primitives not perform as predicted, then the tests for interconnectivity of the two Application Entities shall be suspended.

Each of the Services Classes between the two Application Entities shall be tested individually. Should the testing within a Service Class not perform as predicted, then the testing for that Service Class shall be suspended.

8.2 Resumption Requirements. Testing of the Association Services for the two Application Entities in question will resume from the beginning when the reasons for suspension of Association Services testing have been determined in either Application Entity, have been corrected, and new versions of either or both Application  Entities have been submitted to the testing group.

Testing within a Service Class for the two Application Entities in question will resume from the beginning of testing for that Service Class when the reasons for suspension of testing that Service Class have been determined in either Application Entity and new versions have been submitted to the testing group.

9.	Test Deliverables

The following documents shall be included:

Test Plan

Test Design Specifications

Test Case Specifications

Test Logs

Test Incident Report Logs

Test Summary Reports

Test Input Data and Test Output Data

10.	Testing  Tasks

Identify the set of tasks necessary to prepare for and to perform the tests. Identify all intertask dependencies and any special skills required.

Tasks include:

(1)	Preparation of the test plan

(2)	Preparation of test design specification

(3)	Preparation of each test-case specification

(4)	Preparation of test-procedure specification

(5)	Prepare the hardware test environment

(6)	Prepare the software test environment

(7)	Perform the test procedures

(8)	Resolve test incident reports

(9)	Repeat tasks (6) - (9) until all test procedures are successful

(10)	Prepare the test summary report

(11)	Transmit all test documentation and test data to the appropriate groups

11.	Environment

11.1 Environmental Needs

This information specifies both the necessary and desired properties of the test environment. This specification should contain all the pertinent characteristics of the testing facilities including: the hardware, the communications and system software, the mode of usage (for example, stand-alone), and any other software or facilities needed to support the test.

11.2 Description of Actual Testing Environment

This specification should contain all the pertinent characteristics of the testing facilities including: the hardware, the communications and system software, the mode of usage (for example, stand-alone), and any other software or facilities used to support the test.

12.	Responsibilities

Identify the groups responsible for managing, designing, preparing, executing, witnessing, checking, and resolving all aspects of the testing process. In addition, identify the groups responsible for providing the test items and the environmental needs.

13.	Staffing and Training

Specify test staffing needs by skill level. Identify training options for providing necessary skills.

14.	Schedule

Identify the significant milestones of the testing project.

15.	Approvals

Specify the names and titles of all responsible persons who must approve the test plan. Signatures and dates of approval are desirable.