Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Joint Test Schema Template

Test Plan, Test Design, Test Procedure, Test Cases

Document ID:		1.2.840.113702.1.1.6.1

Version Date:		6/4/96

Application Entities Interconnectivity

Joint Test Schema Template

i.	Purpose.

The purpose of a Test Schema document is to prescribe the scope, approach, resources, and schedule of the testing activities.

To identify the items being tested, the features to be tested, the testing tasks to be performed, the personnel responsible for each task, and the risks associated with the test plan.

This document is the Level_1 Joint Test Schema template document to address complete testing for interconnectivity between two Application Entities from cooperating vendors; it is intended to be the template base document from which other joint test schemata to define such testing are derived. This document is intended to represent a generic Joint Test Schema, capable of testing all of the characteristics of interconnected Application Entities that conform to the DICOM Standard, ACR/NEMA Standards Publication No. PS 3. The format of this joint test schema template is a modification of the documents format defined by the IEEE Testing Standard (IEEE Std. 829-1983). This Test Schema format has been modified to include the essential elements of the Test Plan, Test Design, Test Procedure, and Test Case documents as defined by the IEEE Std. 829-1983.

ii.	Outline.

A test schema document shall have the following structure:

1.	Test schema identifier

2.	Introduction

3.	Test items

4.	Features to be tested

5.	Features not to be tested

6.	Approach

7.	Item pass/fail criteria

8.	Suspension criteria and resumption requirements

9.	Test deliverables

10.	Testing tasks

11.	Environmental needs

12.	Responsibilities

13.	Staffing and training needs

14.	Schedule

15.	Test Unit

15.1	Test Design Specifications

15.2	Test Procedure Specifications

15.3	Test Cases Specifications

16.	Approvals

The Test Design Specifications, Test Procedure Specifications, and Test Cases Specifications form a Test Unit and may be contained directly in the test schema document, in appendices of the test schema document, or they may be contained in separate documents which are referenced at the appropriate points in the test schema document.

The general structure of a test schema must accommodate multilevel testing situations. The highest level test unit of a complex system may utilize a set of lower level test units, each consisting of a test design, test procedure, and test cases. The documents of the lower level test units may be contained in line or referenced at the appropriate points in the upper level  test unit.

1.	Test-Schema Identifier

The Test Schema Identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

2.	Introduction

2.1	Objectives.  A test schema for testing the interconnectivity between two Application Entities shall support the following objectives.

(1)  To detail the activities required to prepare for and support the test.

(2)  To communicate to all responsible parties the tasks which they are to perform and the schedule to be followed in performing the tasks.

(3)  To define the sources of information used to prepare the test plan.

(4)  To define the test tools and environment needed to conduct the test.

(5)	To define the test design, test procedure, and test cases required to conduct the test.

2.2	Background.  The DICOM Standard, ACR/NEMA Standards Publication No. PS 3 has been developed to define a standard method of transferring images and associated information between devices manufactured by different vendors. The standard defines various information objects and the operations or services that may be performed on the objects, organizes the objects and services into Service-Object Pair (SOP) Classes, and groups the SOP Classes into Service Classes which can be supported by any Application Entity. The DICOM Standard also allows each implementor of an Application Entity to determine the degree to which the Application Entity will conform to the DICOM Standard for each Service-Object Pair Class. The standard further distinguishes the services that control the ability to exchange information as the Association Services.

2.3	Scope.  A test schema for interconnectivity between two Application Entities has the objective to show that the Application Entities are able to establish Associations and exchange messages in accordance with the DICOM Standard, ACR/NEMA Standards Publication No. PS 3. A test for interconnectivity between two Application Entities is not intended to be a detailed test process for debugging either or both of the Application Entities involved. Testing should be kept at a level that will demonstrate successful or unsuccessful inter-operation without going down to a level to ‘debug’ why the two Application Entities fail to inter-operate successfully.

The scope of testing  will include functionality testing of the Application Entities, the operator interfaces for invoking message transfer, and indications of successful and failed operations.

This Joint Test Schema presents a framework for organizing the testing of two Application Entities for interconnectivity. The schema partitions the testing process in exactly the same way that the DICOM Standard organizes the Service Classes and the Association Services. The Association Services are distinguished for testing because they are required to be operational before any of the Services Classes can be tested. For the Association Services and each of the supported Service classes, a joint testing cycle for testing interconnectivity between Application Entities of cooperating vendors will be developed, with the accompanying set of documents required to record the design and execution of the testing process.

This Joint Test Schema and any other testing documents at this level serve only to present the framework for organizing testing of the Association Services and the Service Classes. The testing documentation at this level does not attempt to present detailed testing information for the Association Services or any of the Services Classes. That degree of detail is left to the testing documentation associated with the Association Services and each of the Service Classes. The test unit for each category will include functionality testing of the Application Entities, the operator interfaces for invoking message transfer, and indications of successful and failed operations.

2.4	References. The following documents shall be included in the references of every joint test for interconnectivity of two Application Entities:

DICOM Standard, ACR/NEMA Standards Publication No. PS 3

The Conformance Statements of each Application Entity to be tested.

The Test Plan, Test Design, and Test Procedure, and Test Case documents prepared by each vendor for stand-alone testing of their Application Entity.

3.	Test Items

The items to be tested for interconnectivity shall be the two Application Entities of interest. These test items shall be identified down to the level of version/revision number. As specified by the DICOM Standard, the Upper Layer Association Services are included in the Application Entities and are therefore included in the testing process.

The following documents, if available, will be used to provide the definition of correct operation.

DICOM Standard, ACR/NEMA Standards Publication No. PS 3

The Conformance Statements of each Application Entities to be tested.

The Test Plan, Test Design, and Test Procedure, and Test Case documents prepared by each vendor for stand-alone testing of their Application Entity.

4.	Features to be Tested

The following features will be tested in every joint test of interconnectivity between two Application Entities. For any two Application Entities, only the services which are common to both Application Entities can of course be tested.

Association Services

Supported Service Classes and Service Class Roles

Storage Service Class

Query / Retrieve Service Class

Print Service Class

Patient, Study, and Results Management Service Classes

Basic Worklist Management Service Class

Storage Commitment Service Class

Study Content Notification Service Class

New Service Classes as they are incorporated into the Standard

5.	Features Not to be Tested

The scope of the DICOM standard, ACR/NEMA Standards Publication No. PS 3 covers a wide range of features that may not be implemented in both of the Application entities that are to be tested. Features that are not common to both Application Entities will not be tested.

6.	Approach

The test design for a pair of  Application Entities will be organized in a hierarchy of test units.  Testing will begin with the Association Services, followed by tests for each of the Service Classes that are common to the pair of Application Entities. Within  each service category (Association Services, Storage, Query/Retrieve , etc.), a sequence of tests will be performed to test all features that are applicable to the Application Entities. The test teams from the cooperating vendors will use the DICOM Standard, PS 3, the respective Conformance Statements and the test documents of both Application Entities to prepare the joint test schema.

7.	Item Pass/Fail Criteria

The Pass/Fail criterion for interconnection of two Application Entities in the Association Services and each of the Service Classes is that they should exhibit the behavior expressed in their conformance statements, consistent with the DICOM Standard. For each test in each category, the operational events should take place in the proper sequence and each of the Application Entities should be left in the appropriate state following the operation.

8.	Suspension Criteria and Resumption Requirements

8.1 Suspension Criteria. The Association Services are a prerequisite to the operation of all of the Services Classes and must operate successfully in order for testing of any of the Service Classes to proceed. Should any of the Association Services primitives not perform as predicted, then the tests for interconnectivity of the two Application Entities shall be suspended.

Each of the Services Classes jointly supported by the two Application Entities shall be tested individually. Should the testing within a Service Class not perform as predicted, then the testing for that Service Class shall be suspended.

8.2 Resumption Requirements. Testing of the Association Services for the two Application Entities in question will resume from the beginning when the reasons for suspension of Association Services testing have been determined in either Application Entity, have been corrected, and new versions of either or both Application  Entities have been submitted to the testing group.

Testing within a Service Class for the two Application Entities in question will resume from the beginning of the test unit for that Service Class when the reasons for suspension of testing that Service Class have been determined in either Application Entity and new versions have been submitted to the testing group.

9.	Test Deliverables

The following documents shall be included:

Joint Test Schema

Test Logs

Test Incident Report Logs

Joint Test Summary Report

Test Input Data and Test Output Data

10.	Testing  Tasks

Identify the set of tasks necessary to prepare for and to perform the tests. Identify all intertask dependencies and any special skills required.

Tasks include:

(1)	Preparation of the joint test schema specification

(2)	Prepare the hardware test environment

(3)	Prepare the software test environment

(4)	Perform the test procedures

(5)	Resolve test incident reports

(6)	Repeat tasks (3) - (6) until all test procedures are successful

(7)	Prepare the test summary report

(8)	Transmit all test documentation and test data to the appropriate groups

11.	Environment

11.1 Environmental Needs

This information specifies both the necessary and desired properties of the test environment. This specification should contain all the pertinent characteristics of the testing facilities including: the hardware, the communications and system software, the mode of usage (for example, stand-alone), and any other software or facilities needed to support the test.

11.2 Description of Actual Testing Environment

This specification should contain all the pertinent characteristics of the testing facilities including: the hardware, the communications and system software, the mode of usage (for example, stand-alone), and any other software or facilities used to support the test.

12.	Responsibilities

Identify the groups responsible for managing, designing, preparing, executing, witnessing, checking, and resolving all aspects of the testing process. In addition, identify the groups responsible for providing the test items and the environmental needs.

13.	Staffing and Training

Specify test staffing needs by skill level. Identify training options for providing necessary skills.

14.	Schedule

Identify the significant milestones of the testing project.

15.  Test Unit Specification

15.1.  Test Design Specification

The Test Design Specification  may be included in line in the test schema document, in an appendix of the test schema document, or in a separate document. The applicable components from the following structure shall be contained in a Test Design Specification.

15.1.1	Test Design Identifier A Test Design Specification Identifier shall be established for the Test Design Specification. The Test Design identifier shall be comprised of the document identifier with a suffix of the appropriate section within the containing document. No suffix implies that the Test Design Specification is the only information in the containing document. The Test Design identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

15.1.2	Features to be Tested This section will be used to list the intersection of functionality between the two systems under test.

Identify the common Service Classes between the two systems.

Within Service Classes, identify the common Service-Object Pairs (SOPs).

Within the SOPs, identify complementary roles.

Within each service category, all relevant features that are common to both Application Entities will be tested. Implementation specific characteristics and implementation specific policies may be tested.

The inherent hierarchy of complementary roles within SOPs within Service Classes implies that the organization of Test Procedures and Test Cases within a Test Unit is inherently recursive.

15.1.3	Approach Refinements Having identified the common service classes, SOPs, and complementary roles, the next step is to review the tests and procedures established by each vendor. From the tests and procedures established by the vendors, a complementary set must be selected and, perhaps, modified to create the Test Unit for interconnectivity testing.  The Association Services must be operational before any of the Service Classes can be tested, and therefore must be tested first. After the Association Services have been successfully tested, the compatible service classes and SOPs within service classes can be tested in any order established by the testing teams.

15.1.4	Test Identification  List the identifier of each test case associated with this test design and give a brief description of the test case. The test case may be included in line in the Test Schema document, in an appendix of the Test Schema document, or in a separate document. The test case identifier shall be comprised of the document identifier with a suffix of the appropriate section within the containing document. No suffix implies that the test case is the only information in the containing document.

15.1.5	Feature Pass/Fail Criteria   Specify the criteria to be used to determine whether the feature or feature combination has passed or failed.

15.2.  Test Procedure Specification

The Test Procedure Specification  may be included in line in the test schema document, in an appendix of the test schema document, or in a separate document. The applicable components from the following structure shall be contained in each Test Procedure Specification.

15.2.1	Test Procedure Identifier  A Test Procedure Specification Identifier shall be established to denote the steps for executing the test cases and any additional steps that may be required to execute the interconnectivity tests. The Test Procedure identifier shall be comprised of the document identifier with a suffix of the appropriate section within the containing document. No suffix implies that the Test Procedure Specification is the only information in the containing document.The Test Procedure identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

15.2.2	Purpose  Describe the purpose of the test procedure. Reference all of the test cases that the test procedure requires to be executed.

15.2.3	Special Requirements  Identify any special requirements that are necessary for the execution of this test procedure. These special requirements may include prerequisite procedures, special skills requirements, and special environmental requirements.

15.2.4	Procedure Steps  The following steps, where applicable, may be included in the test procedure.

15.2.4.1	Log  Describe any special methods for logging the results of test execution, the incidents observed, and other events pertinent to the test.

15.2.4.2	Set Up  Describe the sequence of actions necessary to prepare for execution of the procedure.

15.2.4.3	Start  Describe the actions necessary to begin execution of the test procedure.

15.2.4.4	Proceed  Describe any actions necessary during execution of the test procedure.

15.2.4.5	Measure  Describe how the test measurements will be made.

15.2.4.6	Shut Down  Describe the actions necessary to suspend testing, when unscheduled events dictate.

15.2.4.8	Restart  Identify any procedural restart points and describe the actions necessary to restart the procedure at each of these points.

15.2.4.9	Stop  Describe the actions necessary to bring execution to an orderly halt.

15.2.4.10	Wrap Up Describe the actions necessary to restore the environment.

15.2.4.11	Contingencies  Describe the actions necessary to deal with anomalous events that may occur during execution.

15.3.  Test Case Specification

For each test case to be executed between the two Application Entities, a Test Case Specification shall be developed. The Test Case Specification  may be included in line in the test schema document, in an appendix of the test schema document, or in a separate document. The applicable components from the following structure shall be contained in each Test Case Specification.

15.3.1	Test Case Identifier The Test Case identifier shall be comprised of the document identifier with a suffix of the appropriate section within the containing document. No suffix implies that the Test Case Specification is the only information in the containing document. The Test Case Specification identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

15.3.2	Test Items  The two items involved are the two Application Entities being tested for interconnectivity.

15.3.3	Input Specifications

Specify each input required to execute the test case, by value, by filename, etc.

Identify all appropriate data bases, files, terminal messages, memory resident areas, etc.

Specify all relationships between inputs.

15.3.4	Output Specifications

Specify all the outputs and features required of the test items.

15.3.5	Environmental Needs

15.3.5.1 Hardware  The computer system required for each of the Application Entities is as defined by the respective suppliers of the Application Entities. The assumption is that the systems are networked together and therefore the appropriate hardware connections to the respective networks are required.

15.3.5.2 Software  The system and application software required for each of the Application Entities is as defined by the respective suppliers of the Application Entities. The defined communication modes for DICOM are: 1) OSI,  2) TCP/IP,  3) Point-to-Point. The assumption is that each system has the appropriate communication software available to the Application Entity.

15.3.5.3 Other environmental requirements  Other environmental requirements that are common to all test cases will be documented in this section of test cases overview. Environmental requirements that are specific to a test case will be documented in the corresponding section of the test case specification.

15.3.6	Special Procedural Requirements

15.3.7	Intercase Dependencies

16.	Approvals

Specify the names and titles of all responsible persons who must approve the test plan. Signatures and dates of approval are desirable.