Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Joint Test SchemaTemplate - Storage

Document ID:		1.2.840.113702.1.2.11.1

Version Date:		06/04/96

Application Entities Interconnectivity

Joint Test Schema Template - Storage

i.	Purpose.

To identify the items being tested, the features to be tested, the testing tasks to be performed, the personnel responsible for each task, and the risks associated with the test plan.

The purpose of a Test Schema document is to prescribe the scope, approach, resources, and schedule of the testing activities.

This document is the Level_2 Joint Test Schema template document to address complete testing for interconnectivity between two Application Entities from cooperating vendors; it is intended to be the template base document from which other joint test schemata are derived to define such testing. This document is intended to represent a generic Joint Test Schema, capable of testing all of the characteristics of interconnected Application Entities that conform to the DICOM Standard, ACR/NEMA Standards Publication No. PS 3. The format of this joint test schema template is a modification of the documents format defined by the IEEE Testing Standard (IEEE Std. 829-1983). The format has been modified such this Test Schema includes the essential elements of the Test Plan, Test Design, Test Procedure, and Test Case documents as defined by the IEEE Std. 829-1983.

ii.	Outline.

A test schema document shall have the following structure:

1.	Test schema identifier

2.	Introduction

3.	Test items

4.	Features to be tested

5.	Features not to be tested

6.	Approach

7.	Item pass/fail criteria

8.	Suspension criteria and resumption requirements

9.	Test deliverables

10.	Testing tasks

11.	Environmental needs

12.	Responsibilities

13.	Staffing and training needs

14.	Schedule

15.	Test Unit

15.1	Test Design Specifications

15.2	Test Procedure Specifications

15.3	Test Cases Specifications

16.	Approvals

1.	Test-Schema Identifier

The Test Schema Identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

2.	Introduction

2.1	Objectives.  A test schema for testing the interconnectivity between two Application Entities shall support the following objectives.

(1)  To detail the activities required to prepare for and support the test.

(2)  To communicate to all responsible parties the tasks to be performed and the schedule to be followed in performing the tasks.

(3)  To define the sources of information used to prepare the test plan.

(4)  To define the test tools and environment needed to conduct the test.

(5)  To define the test procedure required to execute the test plan.

(6)  To define the test cases included in the test design.

2.2	Background.  The DICOM Standard, ACR/NEMA Standards Publication No. PS 3 has been developed to define a standard method of transferring images and associated information between devices manufactured by different vendors. The standard defines various information objects and the operations or services that may be performed on the objects, organizes the objects and services into Service-Object Pair (SOP) Classes, and groups the SOP Classes into Service Classes which can be supported by any Application Entity. The DICOM Standard also allows each implementor of an Application Entity to determine the degree to which the Application Entity will conform to the DICOM Standard for each Service-Object Pair Class. The standard further distinguishes the services that control the ability to exchange information as the Association Services.

2.3	Scope.  A test schema for interconnectivity between two Application Entities has the objective to show that the Application Entities are able to establish Associations and exchange messages in accordance with the DICOM Standard, ACR/NEMA Standards Publication No. PS 3. A test for interconnectivity between two Application Entities is not intended to be a detailed test process for debugging either or both of the Application Entities involved. Testing should be kept at a level that will demonstrate successful or unsuccessful inter-operation without going down to a level to ‘debug’ why the two Application Entities fail to inter-operate successfully.

The scope of testing  will include functionality testing of the Application Entities, the operator interfaces for invoking message transfer, and indications of successful and failed operations.

This Joint Test Schema presents a schema for organizing the testing of two Application Entities for interconnectivity. The schema partitions the testing process in exactly the same way that the DICOM Standard organizes the Service Classes and the Association Services. The Association Services are distinguished for testing because they are required to be operational before any of the Services Classes can be tested. For each of these, a joint testing cycle for testing interconnectivity between Application Entities of cooperating vendors will be developed, with the accompanying set of documents required to record the design and execution of the testing process.

This Joint Test Schema and any other testing documents at this level are templates for organizing testing of the Association Services and the Service Classes.

2.4	References.   The following documents shall be included in the references of every joint test for interconnectivity of two Application Entities:

DICOM Standard, ACR/NEMA Standards Publication No. PS 3

The Conformance Statements of both Application Entities to be tested.

The Test Plan, Test Design, Test Procedure, and Test Cases documents prepared by each vendor for stand-alone testing of their Application Entity.

3.	Test Items

The items to be tested for interconnectivity shall be the two Application Entities of interest. These test items shall be identified down to the level of version/revision number. As specified by the DICOM Standard, the Upper Layer Association Services are included in the Application Entities and are therefore included in the testing process.

The following documents, if available, will be used to provide the definition of correct operation.

DICOM Standard, ACR/NEMA Standards Publication No. PS 3

The Conformance Statements of each Application Entities to be tested.

The Test Plan, Test Design, and Test Procedure, and Test Case documents prepared by each vendor for stand-alone testing of their Application Entity.

4.	Features to be Tested

The following features will be tested in every joint test of interconnectivity between two Application Entities. For any two Application Entities, only the services which are common to both Application Entities can of course be tested.

Association Services

Supported Service Classes and Service Class Roles

Storage Service Class

Query / Retrieve Service Class

Print Service Class

Patient, Study, and Results Management Service Classes

Basic Worklist Management Service Class

Storage Commitment Service Class

Study Content Notification Service Class

New Service Classes as they are incorporated into the Standard

Within each service category, all relevant features that are common to both Application Entities will be tested.

5.	Features Not to be Tested

The scope of the DICOM standard, ACR/NEMA Standards Publication No. PS 3 covers a wide range of features that may not be implemented in both of the Application entities that are to be tested. Features that are not common to both Application Entities will not be tested.

6.	Approach

The test design for a pair of  Application Entities will be organized in a hierarchy of testing units.  Before testing can proceed for the Storage Service Class, the Association Services must have already been tested or must be tested. Within  each service category, in this case the Association Services and Storage Service Class, a sequence of tests will be performed to test all features which are applicable to the pair of Application Entities. The test teams from the cooperating vendors will use the DICOM Standard, PS 3, the respective Conformance Statements and documentation of the Application Entities to prepare the joint test schema for the two service categories.

7.	Item Pass/Fail Criteria

The Pass/Fail criterion for interconnection of two Application Entities in each of the Service Classes and the Association Services is that they should exhibit the behavior expressed in their conformance statements, consistent with the DICOM Standard. For each test in each category, the operational events should take place in the proper sequence and each of the Application Entities should be left in the appropriate state following the operation.

8.	Suspension Criteria and Resumption Requirements

8.1	Suspension Criteria. The Association Services are a prerequisite to the operation of all of the Services Classes and must operate successfully in order for testing of any of the Service Classes to proceed. Should any of the Association Services primitives not perform as predicted, then the tests for interconnectivity of the two Application Entities shall be suspended.

Each of the Services Classes jointly supported by the two Application Entities shall be tested individually. Should the testing within a Service Class not perform as predicted, then the testing for that Service Class shall be suspended.

8.2	Resumption Requirements. Testing of the Association Services for the two Application Entities in question will resume from the beginning when the reasons for suspension of Association Services testing have been determined in either Application Entity, have been corrected, and new versions of either or both Application  Entities have been submitted to the testing group.

Testing within a Service Class for the two Application Entities in question will resume from the beginning of the test unit for that Service Class when the reasons for suspension of testing that Service Class have been determined in either Application Entity, have been corrected, and new versions have been submitted to the testing group.

9.	Test Deliverables

The following documents shall be included:

Joint Test Schema

Test Logs

Test Incident Report Logs

Joint Test Summary Report

Test Input Data and Test Output Data

10.	Testing  Tasks

Identify the set of tasks necessary to prepare for and to perform the tests. Identify all intertask dependencies and any special skills required.

Tasks include:

(1)	Preparation of the joint test schema

(2)	Prepare the hardware test environment

(3)	Prepare the software test environment

(4)	Perform the test procedures

(5)	Resolve test incident reports

(6)	Repeat tasks (3) - (6) until all test procedures are successful

(7)	Prepare the test summary report

(8)	Transmit all test documentation and test data to the appropriate groups

11.	Environment

11.1	Environmental Needs

This information specifies both the necessary and desired properties of the test environment. This specification should contain all the pertinent characteristics of the testing facilities including: the hardware, the communications and system software, the mode of usage (for example, stand-alone), and any other software or facilities needed to support the test.

11.2	Description of Actual Testing Environment

This specification should contain all the pertinent characteristics of the testing facilities including: the hardware, the communications and system software, the mode of usage (for example, stand-alone), and any other software or facilities used to support the test.

12.	Responsibilities

Identify the groups responsible for managing, designing, preparing, executing, witnessing, checking, and resolving all aspects of the testing process. In addition, identify the groups responsible for providing the test items and the environmental needs.

13.	Staffing and Training

Specify test staffing needs by skill level. Identify training options for providing necessary skills.

14.	Schedule

Identify the significant milestones of the testing project.

15.  Test Unit Specification

15.1.  Test Design Specification

In this case, the Test Design Specification is contained within the Test Schema document.

15.1.1	Test Design Identifier

15.1.2	Features to be Tested The following features will be tested for Storage Class interconnectivity between two Application Entities. All composite SOPs and complementary roles that the two Application Entities mutually support must be tested.

Association Services:

A-ASSOCIATE	Establish an Association

A-RELEASE	Release an Association

A-ABORT	Abort an Association

C-ECHO	Test for DICOM connectivity

P-DATA	Send a DICOM message

Storage Service:

C-STORE	Store a composite SOP Instance from the Service Class User (SCU) to the Service Class Provider(SCP).

15.1.3	Approach Refinements Interconnectivity testing of two Application Entities for the Storage Service Class will be organized in a hierarch of test sequences. The Association Services must already have been tested or must be tested before testing the Storage Service Class can proceed. Within the Storage Service Class, all SOPs and complementary roles that are mutually supported by the two AE’s must be tested in sequence.

15.1.4	Test Identification   Each test to be performed shall be identified in this test design specification. For this template test schema, the Association Services and the Storage Service Class are the categories to be tested.

Association Services:

A-ASSOCIATE	Establish an Association		[Appendix A.1 ]

A-RELEASE	Release an Association		[Appendix A.2 ]

A-ABORT	Abort an Association		[Appendix A.3 ]

C-ECHO	Test for DICOM connectivity	[Appendix A.4 ]

P-DATA	Send a DICOM message		[Appendix A.5 ]

Storage Service:

C-STORE	Store a composite SOP Instance from the Service Class User (SCU) to the Service Class Provider(SCP).	[Appendix B.1 ]

15.1.5	Feature Pass/Fail Criteria   The Pass/Fail criteria for testing the Association Services at this level are expressed in terms of the events that can be monitored and that can be expected to be logged in the log files of the Application Entities and the Upper Layer Association Services.

15.2.  Test Procedure Specification

In this case, the Test Procedure Specification is contained within the test schema document.

15.2.1	Test Procedure Identifier

15.2.2	Purpose  The purpose of this test procedure is to test the Association Services and all Service Classes that are common to the pair of  Application Entities of interest. Tests will be performed for each of the following features:

Association Services:

A-ASSOCIATE	Establish an Association		[Appendix A.1]

A-RELEASE	Release an Association		[Appendix A.2 ]

A-ABORT	Abort an Association		[Appendix A.3 ]

C-ECHO	Test for DICOM connectivity	[Appendix A.4 ]

P-DATA	Send a DICOM message		[Appendix A.5 ]

Storage Service:

C-STORE	Store a composite SOP Instance from the Service Class User (SCU) to the Service Class Provider(SCP).	[Appendix B.1]

15.2.3	Special Requirements  The procedure to be followed in executing the tests is independent of the physical palcement of the Application Entities, which may be in the same location or very different locations. Each Application Entity must have an operator / observer and the operators must have a reliable means of communication in addition to the network connection of the Application Entities. For each test to be performed, the operator at the initiator Application Entity (Service Class User, Service Class Provider) will initiate the operation. At both Application Entities, the operator / observers will track the sequence of events that occur. Comparing the observed sequence of events to the sequence anticipated from the conformance statements will determine if the Application Entities are interacting as predicted.

15.2.4	Procedure Steps  Interconnectivity testing of a pair of Application Entities will be conducted as a hierarchy of test sequences. For the Storage Service Class, the upper level sequence involves the Association Services followed by the Storage Service Class. Within each service category, operations will be tested in sequence to ensure that they perform as expected.

Association Services:

A-ASSOCIATE	Establish an Association

A-RELEASE	Release an Association

A-ABORT	Abort an Association

C-ECHO	Test for DICOM connectivity

P-DATA	Send a DICOM message

Storage Service:

C-STORE	Store a composite SOP Instance from the Service Class User (SCU) to the Service Class Provider(SCP).

15.2.4.1	Log  Logging the input and output sequence of events at both Application Entities is required. At each Application Entity, it is essential to tell if it is receiving the expected inputs and is putting out the predicted outputs.

15.2.4.2	Set Up  Ensure that network connectivity has been established for the Application Entity pair and that both are equipped for normal operation.

15.2.4.3	Start  Describe the actions necessary to begin execution of the test procedure.

15.2.4.4	Proceed  Describe any actions necessary during execution of the test procedure.

15.2.4.5	Measure  Describe how the test measurements will be made.

15.2.4.6	Shut Down  Describe the actions necessary to suspend testing, when unscheduled events dictate.

15.2.4.8	Restart  Identify any procedural restart points and describe the actions necessary to restart the procedure at each of these points.

15.2.4.9	Stop  Describe the actions necessary to bring execution to an orderly halt.

15.2.4.10	Wrap Up Describe the actions necessary to restore the environment.

15.2.4.11	Contingencies  Describe the actions necessaryto deal with anomolous events that may occur during execution.

15.3.  Test Case Specifications

In this case, the Test Case Specifications are contained in Appendices of this test schema document. This section contains the outline form used for each test case specification.

15.3.1	Test Case Identifier

15.3.2	Test Items  The two items involved are the two Application Entities being tested for interconnectivity.

15.3.3	Input Specifications

Specify each input required to execute the test case, by value, by filename, etc.

Identify all appropriate data bases, files, terminal messages, memory resident areas, etc.

Specify all relationships between inputs.

15.3.4	Output Specifications

Specify all the outputs and features required of the test items.

15.3.5	Environmental Needs

15.3.5.1 Hardware  The computer system required for each of the Application Entities is as defined by the respective suppliers of the Application Entities.

The assumption is that the systems are networked together and therefore the appropriate hardware connections to the respective networks are required.

15.3.5.2 Software  The system and application software required for each of the Application Entities is as defined by the respective suppliers of the Application Entities.

The defined communication modes for DICOM are: 1) OSI,  2) TCP/IP,  3) Point-to-Point. The assumption is that each system has the appropriate communication software available to the Application Entiry.

15.3.5.3 Other environmental requirements  Other environmental requirements that are common to all test cases will be documented in this section of test cases overview. Environmental requirements that are specific to a test case will be documented in the corresponding section of the test case specification.

15.3.6	Special Procedural Requirements

15.3.7	Intercase Dependencies

16.	Approvals

Specify the names and titles of all responsible persons who must approve the test plan. Signatures and dates of approval are desirable.

Appendix A

Association Services  -  Test Cases

Association Services:

A-ASSOCIATE		Establish an Association

A-RELEASE		Release an Association

A-ABORT		Abort an Association

C-ECHO		Test for connectivity

P-DATA		Send a DICOM message

Association Service Provider Internal Integrity:

A-P-ABORT

A.1	A-ASSOCIATE:

A.1.1	Test Case Identifier:

The test-case-specification identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| A-ASSOCIATE | Establish an Association | Specified by testing team |

A.1.2	Test Items:

The test items involved are the two Application Entities being tested for interconnectivity.

A.1.3	Input Specifications:

The input specifications required for testing the A-ASSOCIATE Upper Layer service shall be the user interaction required to cause the requestor Application Entity to issue an A-ASSOCIATE request.

Documented as part of the input specification for the test of the A-ASSOCIATE service primitive shall be the following items:

Application Context

Presentation Contexts

User Information Items

Two user interaction input steps are required at the requestor Application Entity to:

1.	cause the A-ASSOCIATE Upper Layer service to be invoked, and

2.	cause the C-ECHO DIMSE service to be invoked to confirm that the A-ASSOCIATE service has been successful.

A.1.4	Output Specifications:

The output and features expected of the Application Entities for successful execution of the A-ASSOCIATE service are as follows:

Output:

Log files entries:

Event Level Information:

Association Create

| Requestor | Service Primitive Direction |  |  | Acceptor |
| --- | --- | --- | --- | --- |
| A-ASSOCIATE request |  |  |  |  |
|  |  |  | A-ASSOCIATE indication |  |
|  |  |  | A-ASSOCIATE response |  |
| A-ASSOCIATE confirmation |  |  |  |  |
| Note: At this point, the Association request may have been accepted of rejected. If the Association has been accepted, then either Application Entity may test for connectivity at the application level by issuing the DIMSE C-ECHO service primitive. |  |  |  |  |
| C-ECHO request |  |  |  |  |
|  |  |  | C-ECHO indication |  |
|  |  |  | C-ECHO response |  |
| C-ECHO confirmation |  |  |  |  |

Data Level Information:

The data recorded from testing the A-ASSOCIATE service primitive should include whether the Association was Accepted or Rejected, the Presentation Contexts accepted or rejected, the Transfer Syntax accepted for each accepted Presentation Context, and the agreed upon user items.

Features:

At this point, the association between the two Application Entities should be established and ready for message transfer.

A.1.5	Environmental Needs:

A.1.6	Special Procedural Requirements:

A.1.7	Intercase Dependencies:

The A-ASSOCIATE test caae must be executed prior to executing any of the other Association Services test cases.

A.2	A-RELEASE

A.2.1	Test Case Identifier

The test-case-specification identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| A-RELEASE | Release an Association | Specified by testing team |

A.2.2	Test Items

The test items involved are the two Application Entities being tested for interconnectivity.

A.2.3	Input Specifications

The input specifications required for testing the A-RELEASE Upper Layer service shall be the user interaction required to cause the requestor Application Entity to issue an A-RELEASE request.

Depending on the level of control afforded to an operator by the designers and implementers, the testers may also be able to explicitly test for the case of  ‘Association release service procedure collision’.

The assumption in this test is that the A-ASSOCIATE service has previously been successfully invoked.

Output:

A.2.4	Output Specifications

The output and features expected of the Application Entities for successful execution of the A-RELEASE service are as follows:

Log files entries:

Event Level Information:

Normal Association Release

| Requestor | Service Primitive Direction |  |  | Acceptor |
| --- | --- | --- | --- | --- |
| A-RELEASE request |  |  |  |  |
|  |  |  | A-RELEASE indication |  |
|  |  |  | A-RELEASE response |  |
| A-RELEASE confirmation |  |  |  |  |

Association Release Service Collision

| Association Requestor | Service Primitive Direction |  |  | Association Acceptor |
| --- | --- | --- | --- | --- |
| A-RELEASE request |  |  | A-RELEASE request |  |
| A-RELEASE indication |  |  | A-RELEASE indication |  |
| A-RELEASE response |  |  |  |  |
|  |  |  | A-RELEASE confirmation |  |
|  |  |  | A-RELEASE response |  |
| A-RELEASE confirmation |  |  |  |  |

Features:

At this point, the association between the two Application Entities has been released and no further communication is allowed.

A.2.5	Environmental Needs:

A.2.6	Special Procedural Requirements:

A.2.7	Intercase Dependencies:

The A-ASSOCIATE service must be tested successfully before the A-RELEASE service can be tested.

In performing the test for the A-RELEASE service, the A-ASSOCIATE must be successfully invoked before the A-RELEASE service can be invoked.

A.3	A-ABORT

A.3.1	Test Case Identifier

The test-case-specification identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| A-ABORT | Abort an Association | Specified by testing team |

A.3.2	Test Items

The test items involved are the two Application Entities being tested for interconnectivity.

A.3.3	Input Specifications

The input specifications required for testing the A-ABORT Upper Layer service shall be the user interaction required to cause the requestor Application Entity to issue an A-ABORT request.

The assumptions for conducting this test are as follows:

1.	the ability to invoke the A-ABORT service is accessible to the operator

2.	the A-ASSOCIATE service has previously been successfully invoked.

A.3.4	Output Specifications

The output and features expected of the Application Entities for successful execution of the A-ABORT service are as follows:

Output:

Log files entries:

Event Level Information:

| Requestor | Service Primitive Direction |  | Acceptor |
| --- | --- | --- | --- |
| A-ABORT request |  |  |  |
|  |  |  | A-ABORT indication |

Features:

At this point, the association between the two Application Entities has been released and no further communication is allowed.

A.3.5	Environmental Needs

A.3.6	Special Procedural Requirements

A.3.7	Intercase Dependencies

The A-ASSOCIATE service must be tested successfully before the A-ABORT service can be tested.

In performing the test for the A-ABORT service, the A-ASSOCIATE must be successfully invoked before the A-ABORT service can be invoked.

A.4	C-ECHO

A.4.1	Test Case Identifier

The test-case-specification identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| C-ECHO | Test for connectivity at the Application Entity Level | Specified by testing team |

A.4.2	Test Items

The test items involved are the two Application Entities being tested for interconnectivity.

A.4.3	Input Specifications

The input specifications required for testing the C-ECHO Upper Layer service shall be the user interaction required to cause the requester Application Entity to issue an C-ECHO request.

The assumptions for conducting this test are as follows:

1.	the ability to invoke the C-ECHO service is accessible to the operator

2.	the A-ASSOCIATE service has previously been successfully invoked.

A.4.4	Output Specifications

The output and features expected of the Application Entities for successful execution of the C-ECHO service are as follows:

Output:

Log files entries:

Event Level Information:

| Requester | Service Primitive Direction |  | Acceptor |
| --- | --- | --- | --- |
| C-ECHO request |  |  |  |
|  |  |  | C-ECHO indication |
|  |  |  | C-ECHO response |
| C-ECHO confirmation |  |  |  |

Features:

At this point, Association connectivity between the two Application Entities has been confirmed.

A.4.5	Environmental Needs

A.4.6	Special Procedural Requirements

A.4.7	Intercase Dependencies

The A-ASSOCIATE service must be tested successfully before the C-ECHO DIMSE service can be invoked.

A.5	P-DATA

A.5.1	Test Case Identifier

The test-case-specification identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| P-DATA | Send a DICOM message | Specified by testing team |

A.5.2	Test Items

The test items involved are the two Application Entities being tested for interconnectivity.

A.5.3	Input Specifications

The input specifications required for testing the P-DATA Upper Layer service shall be the user interaction required to cause the requestor Application Entity to issue an P-DATA request.

The assumptions for conducting this test are as follows:

1.	the ability to invoke the P-DATA service is accessible to the operator

2.	the A-ASSOCIATE service has previously been successfully invoked.

A.5.4	Output Specifications

The output and features expected of the Application Entities for successful execution of the P-DATA service are as follows:

Output:

Log files entries:

Event Level Information:

| Requestor | Service Primitive Direction |  |  | Acceptor |
| --- | --- | --- | --- | --- |
| P-DATA request |  |  |  |  |
|  |  |  | P-DATA indication |  |

Features:

At this point, a copy of the DICOM message sent by the requester should be available at the acceptor Application Entity.

A.5.5	Environmental Needs

A.5.6	Special Procedural Requirements

A.5.7	Intercase Dependencies

The A-ASSOCIATE service must be tested successfully before the P-DATA service can be invoked.

Appendix B

Storage Service Class  -  Test Cases

Storage Service Class:

C-STORE		Composite Object Storage

B.1	C-STORE

B.1.1	Test Case Identifier

The test-case-specification identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| C-STORE; Service Object Pair | Storage Service for Service Object Pair | Specified by testing team |

B.1.2	Test Items

The test items involved are the two Application Entities being tested for inter-operability.

B.1.3	Input Specifications

The input specifications required for testing the DIMSE C-STORE service shall be the user interaction required to cause the Service Class User Application Entity to issue a C-STORE request for the Service Object Pair under test.

The user interaction input step at the Service Class User Application Entity is that required to:

1.	cause the Application Entity to issue a DIMSE C-STORE  request for the Service Object Pair under test.

B.1.4	Output Specifications

The output and features expected of the Application Entities for successful execution of the C-STORE service fall into three lavels or categories:

1.	Command Event Level

2.	Command Data Level

3.	Data Set Level

Output:

Log files entries:

Event Level Information:

| Service Class User | Service Primitive Direction |  | Service Class Provider |
| --- | --- | --- | --- |
| C-STORE request |  |  |  |
|  |  |  | C-STORE indication |
|  |  |  | C-STORE response |
| C-STORE confirmation |  |  |  |

Command Data Level Information:

The critical output information at the Command Data Level is the status information sent back to the Service Class User Application Entity in which the Service Class Provider Application Entity  indicates the result of the service. The C-STORE status choices, defined by the DICOM Standard PS3, are given in the following table.

| C-STORE Status |  |
| --- | --- |
| Service Status | Further Meaning |
| Refused | Out of Resources |
|  | Data Set does not match SOP Class |
| Error | Cannot Understand |
|  | Coercion of Data Elements |
| Warning | Data Set does not match SOP Class |
|  | Elements discarded |
| Success |  |

Data Set Level Information:

Features:

B.1.5	Environmental Needs

B.1.6	Special Procedural Requirements

B.1.7	Intercase Dependencies

An Association must have been created and in place in order to perform the C-STORE operation.