Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level_4  Test Summary Report Example

Document ID:		1.2.840.113702.1.4.4.1

Version Date:		11/12/96

Applications Entities Interconnectivity

Level_4 Test Summary Report Example

i.	Purpose

To summarize the results of interconnectivity testing between a Picker PQ2000 CT scanner and a General Electric Advantage Review Workstation and to provide evaluations based on these results.

ii.	Outline

A test-summary report shall have the following structure:

(1)	Test-summary report identifier

(2)	Summary

(3)	Variances

(4)	Comprehensive assessment

(5)	Summary of results

(6)	Evaluation

(7)	Summary of activities

(8)	Approvals

1.	Test-Summary Report

The Test Summary Identifier shall be a unique identifier for both suppliers of the Application Entities tested. This Test Summary Example has the ID: 1.2.840.113702.1.4.4.1.

2.	Summary

The items tested for interconnectivity were the Local DICOM Import Application Entity as implemented in Software Revision 1.03 of the General Electric Advantage Review Workstation and  the HANI Application Entity as implemented in Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner.

A GE Advantage Review Workstation and Picker PQ2000 CT scanner were temporarily taken out of service at the Milton S. Hershey Medical Center of Penn State and dedicated for testing.

The following features were tested:

Association Services

CT Image Information Object Storage SOP Class (UID: 		1.2.840.10008.5.1.4.1.1.2)

The HANI AE served as the SCU and the Local DICOM Import AE served as the SCP.

A Joint Test Schema (ID: 1.2.840.113702.1.4.1.1) was executed on November 3-4, 1995 by Fred Prior, Ph.D. and Franklyn Bradshaw, Ph.D.  Test logs were generated for each test item:

Test Log ID: 1.2.840.113702.1.4.2.1  - HANI Application Entity

Test Log ID: 1.2.840.113702.1.4.3.1  - Local DICOM Import Application Entity.

One Test Incident Report was also generated (ID: 1.2.840.113702.1.4.5.1).

3.	Variances

No  variances of the test items from their conformance statements were encountered. No variances from the Joint Test Schema were introduced..

4.	Comprehensiveness Assessment

Basic interconnectivity was adequately tested.  It was not possible to test the structure or content of the exchanged messages.

5.	Summary of results

Of the nine test cases included in the Joint Test Schema the SCU (HANI) passed 7 unequivocally and the SCP (Local DICOM Import) passed 6.  The Test Incident Report details the anomalous behavior exhibited in the remaining cases.

6.	Evaluation

The HANI AE responded correctly to all test cases, however, in the resource boundary (disk full) and repeatability tests the response was not what had been anticipated.  This appears to be due to the behavior of the SCP.  When instructed to abort an association (the Association Abort test case) the HANI AE appeared to comply reporting an abort after 2 images transferred. It apparently continued to transfer an additional 13 images before actually aborting the association.

In the repeatability of transfers test the Local DICOM Import AE accepted transfer of the same image set twice, returned no warning, and stored both studies in its local directory.  With its storage devices essentially full (too little storage remaining to store an additional image) the Local DICOM Import AE accepted the transfer of a complete image set (a study) and returned a success status.  The images were subsequently lost during a failed attempt at internal format conversion.  Although DICOM does not require persistent storage, this behavior violates the intent of the storage SOP class.

7.	Summary of Activities

Using the Joint Test Schema and its associated test cases the following interconnectivity features were tested:

Basic TCP/IP connectivity

Basic Association Establishment and image transfer

SCU role of the CT Image Storage SOP Class (Picker)

SCP role of the CT Image Storage SOP Class (General Electric)

Association Establishment Policies (Picker):

Single Association Mode

Multiple Association Mode

Association Abort

Behavior at Boundary Cases

Repeatability

The tests required approximately 6 hours spread over two days.  System configuration in preparation for testing required 16 hours.

8.	Approvals

NEMA MEDPACS Contract with Information and Control Technology Inc. -  To Develop Guidelines for Cross-Vendor DICOM Testing, August 2, 1995.